## Celery
Allergen type | Citing Papers | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target | Is method commercially available, or does it have sufficient validation data? | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredients | List relevant/ pertinent untested matrices | Is there more than one extraction method dependant on matrix? | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data. | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Celery | - | SureFood Allergen Celery, Ref. no. S3605 | DNA | Commercially available | declares mg of celery per kg food sample | Real-time PCR | Quantitative | mg celery/kg food | Not detailed in instructions | ≤0.4 mg/kg, determined by manufacturer in corn flour. LOD is dependant on matrix. | 1 mg/kg, determined by manufacturer in corn flour. LOQ is dependant on matrix. | Corn flour | Only corn flour I detailed in the instructions | None detailed in manual. Assay interferents listed as salt, spices, herbs, algae, cocoa 'and further samples matrices' listed as having PCR-inhibiting effects in instructions. | Data too sparce to comment. | Yes, SureFood PREP advanced or SureFast Mag PREP protocols. | Not detailed in kit instructions. | Not detailed in kit instructions. | Not detailed in kit instructions. | Not detailed in kit instructions. | Not detailed in kit instructions. | Data too sparce to comment. | Only method currently used in Fapas® PT trials for celery. More validation data required in manual.

## Cereals containing gluten
Allergen type | Citing Papers | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredient | List relevant/pertinent untested matrice | Is there more than one extraction method dependant on matrix? | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data (inc. precision and recovery). | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Gliadin | - | Rbiopharm RIDASCREEN FAST gliadin sensitive R7051 | R5 antibody plus further antibodies against other gluten proteins | Report as ppm gluten | ELISA | Quantitative | mg/kg | Conversion of gliadin to gluten. Gliadin levels in rye and barley may be over-estimated compared to wheat. | 0.2 mg / kg (ppm) Gliadin (mean) or to 0.4 mg / kg (ppm) Gluten 0.19 – 0.21 mg / kg (ppm) Gliadin* *depending on matrix | 1.25 mg / kg (ppm) Gliadin or 2.5 mg / kg (ppm) Gluten | Gluten containing cereals in as gluten-free declared food (cookies, rice flour, oat flour, sausages, quinoa flour, chocolate, millet flour) | Cookies, rice flour, oat flour, sausages, quinoa flour, chocolate, millet flour | Our interpretation of the manual suggests that cross-reactivities have been determined in raw foods but only on one processed food. Data not included in manual. | Manual not clear regarding scope of testing on processed foods. Raw foods tested.Insufficient data to form an opinion. | Different sample extractions validated depending on the Matrices (Cocktail patented (R7016), Cocktail ECO (R7080)) adapted to the specific requirements of the different samples, methods specified for liquid food samples, other food samples (eg soy and quinoa containing), tannin and polyphenol containing, meat and sausages, and oat samples. Strong acidic or alkaline samples have to be neutralized prior to analysis | Cocktail (patented) (R7006 / R7016) is the official R5-Mendez extraction method according to the Codex Alimentarius and the AOAC | Yes | None identified | Data not specified in manual | Data not specified in manual | In a composed / processed food (e.g. maize bread) cross reactivities might be different. Interfering substances (e.g. polyphenols) can be detected by spike experiments. Quantification may vary due to extraction method used. Only one matrix tested for cross-reactivity. | Validation data and applicabilty data for processed matrices not in manual.
Gliadin | - | Rbiopharm RIDASCREEN FAST gliadin R7002 | Monoclonal antibody R5 reacts with the gliadin- fractions from wheat and corresponding prolamins from rye and barley | Yes | ELISA | Quantitative | mg gliadin /kg | Conversion of gliadin to gluten. Gliadin levels in rye and barley may be over-estimated compared to wheat. | Cocktail (patented) 0.50 mg / kg (ppm) Gliadin (mean) or to 1.0 mg / kg (ppm) Gluten Cocktail ECO 0.52 mg / kg (ppm) Gliadin (mean) or to 1.04 mg / kg (ppm) Gluten (depending on matrix) | 5 mg / kg (ppm) Gliadin or 10 mg / kg (ppm) Gluten | Contaminations by prolamins from wheat (gliadin), rye (secalin), and barley (hordein) in raw products like flours (buckwheat, rice, corn, oats, teff) and spices as well as in processed food like noodles, ready-to-serve meals, bakery products, sausages, beverages and ice cream | Wheat (gliadin), rye (secalin), and barley (hordein) in raw products like flours (buckwheat, rice, corn, oats, teff) and spices as well as in processed food like noodles, ready-to-serve meals, bakery products, sausages, beverages and ice cream | Our interpretation of the manual suggests that cross-reactivities have been determined in raw foods but only on one processed food. Data not included in manual. | Sour dough, beer | Different sample extractions validated depending on the Matrices (Cocktail patented (R7016), Cocktail ECO (R7080)) adapted to the specific requirements of the different samples, methods specified for liquid food samples, other food samples (eg soy and quinoa containing), tannin and polyphenol containing, meat and sausages, and oat samples. Strong acidic or alkaline samples have to be neutralized prior to analysis | Cocktail (patented) (R7006 / R7016) is the official R5-Mendez method according to the Codex Alimentarius and the AOAC | Not stated in manual. | None identified | Data not specified in manual, recovery may be reduced in sourdough or beer. | Data not specified in manual | In a composed / processed food (e.g. maize bread) cross reactivities might be different. Interfering substances (e.g. polyphenols) can be detected by spike experiments | Sensitive method but performance data not included in manual. Performance data for processed food would be beneficial to form an assessment.
Gliadin | van den Broeck, H. C., et al. (2009). "A modified extraction protocol enables detection and quantification of celiac disease-related gluten proteins from wheat." Journal of Chromatography B 877(10): 975-982. | Rbiopharm RIDASCREEN Gliadin competitive R7021 | Monoclonal R5-antibodies | Yes | ELISA | Quantitative | mg gliadin /kg | Conversion of gliadin to gluten. Gliadin levels in rye and barley may be over-estimated compared to wheat. | 2.3 mg/kg (ppm) Gliadin* or to 4.6 mg/kg (ppm) Gluten* 1.9 - 2.6 mg/kg (ppm) Gliadin *mean value | 5 mg/kg (ppm) Gliadin or to 10 mg/kg (ppm) Gluten | Not detailed in manual | Beer, starch syrup, starch, malt extract, sourdough, soy sauce | None detailed in the manual. | Data in manual too sparce to evaluate | Methods specified for solid food (eg starch and starch syrup), polyphenol-containing solid food (e.g. malt and hops), liquid food (e.g. soy sauce), and beer. Strong acidic or alkaline samples have to be neutralized prior to analysis | The competitive ELISA is an AACC international approved method (38-55.01) and an AOAC approved Official Method of Analysis (Final Action OMA 2015.05). The competitive ELISA is an ASBC international approved method (Beer-49). | Yes | The ELISA RIDASCREEN® Gliadin competitive was evaluated by 18 labs in an international collaborative study (AACCI and WGPAT) for beer, starch syrup and sourdough. In a second international collaborative study (ASBC) RIDASCREEN® Gliadin competitive was evaluated by 15 labs for 5 different beers. This is an AOAC approved Official Method of Analysis (Final Action OMA 2015.05) | Recovery only detailed for two matrices at two levels, 103-133%. Repeatability and reproducibility vary greatly between matrices, highlighting the need for a wide range of incurred reference materials to calibrate methods to. | Data not specified in manual | In a composed / processed food (e.g. maize bread) cross reactivities might be different. Interfering substances (e.g. polyphenols) can be detected by spike experiments | AOAC and AACC approved method.
Gliadin | Sharma, G. M. (2012). "Immunoreactivity and Detection of Wheat Proteins by Commercial ELISA Kits." Journal of Aoac International 95(2): 364-371.(Under reported wheat flour levels (74%)) Sharma, G. M., et al. (2013). "Development of an Incurred Cornbread Model for Gluten Detection by Immunoassays." Journal of Agricultural and Food Chemistry 61(49): 12146-12154. Recoveries R-Biopharm (61−108%) -( greater recoveries for gluten vs wheat flour) Gomaa, A. and J. I. Boye (2013). "Impact of thermal processing time and cookie size on the detection of casein, egg, gluten and soy allergens in food." Food Research International 52(1): 483-489.Amnuaycheewa, P., et al. (2022). "Challenges in Gluten Analysis: A Comparison of Four Commercial Sandwich ELISA Kits." Foods 11(5). | Rbiopharm RIDASCREEN Gliadin R7001 | Monoclonal antibody R5 reacts with the gliadin- fractions from wheat and corresponding prolamins from rye and barley | Yes | ELISA | Quantitative | mg gliadin/kg | Conversion of gliadin to gluten. Gliadin levels in rye and barley may be over-estimated compared to wheat. | 0.5 mg / kg (ppm) Gliadin (mean) or 1 mg / kg (ppm) Gluten 0.06 - 1.24 mg / kg (ppm) Gliadin* (*depending on matrix) | 2.5 mg / kg (ppm) Gliadin or 5 mg / kg (ppm) Gluten | Not detailed in manual | Processed and unprocessed rice- and corn-based products, soy, starches, pseudo cereals, legumes, spices, juice, nut nougat crème, cream cheese, pesto, meat, vegetarian meat alternative, cookies, dessert, cake, fish, bread, candies, and potatoes | Not detailed in manual | Soup | Different sample extractions validated depending on the Matrices (Cocktail patented (R7016), Cocktail ECO (R7080)) adapted to the specific requirements of the different samples, methods specified for liquid food samples, other food samples (eg soy and quinoa containing), tannin and polyphenol containing, meat and sausages, and oat samples. Strong acidic or alkaline samples have to be neutralized prior to analysis | RIDASCREEN® Gliadin received from the AOAC as first assay worldwide the status as Official Method of Analysis (AOAC OMA 2012.01) “in food” in 2021. A limitation on certain validated food matrices as usual for other AOAC approvals is not present any longer. | Not detailed in manual but this method uses the official AOAC extraction method. | The publication of the 3rd international collaborative proficiency study from 2020 describes the preparation of 900 samples of 19 differently processed foods from 16 food categories and the obtained results demonstrating it suitability for gluten detection in gluten-free foods. Also compared by Shall et al 2020 to Romer AgraQuant Gluten G12 Assay COKAL02002 kit to determine gluten in wheat. Results anonymised. Both kits showed approx. 100% recovery and resulted varied significantly depending on wheat cultivar. | Not detailed in manual | Sharma et al 2012 LOQ aligned with that stated in manual. Reacted strongly with gliadin fraction rather than glutenenin fraction. Recovery 74% (spiked flour). Responded favourably comapred to other kits. Sharma et al 2013. Good precision. Recovery 61-108%. Recovery reduces with processing. Gave lower recovery compared to other kits for wheat flour compared to gluten-incurred corn bread. Amnuaycheewa et al 2022 Responded similarly to the other R5 and G12 methods in the study, although reported differently for yeast extract, hemp protein powder and cookie. Despite using a different anttibody, responded similarly to a G12 kit in a study to determine wheat cultivars with 100% recoveries for certain cultivars. | In a composed / processed food (e.g. maize bread) cross reactivities might be different. Interfering substances (e.g. polyphenols) can be detected by spike experiments Sample preparation for polyphenol containing raw materials (e.g. chocolate, coffee, cacao, buckwheat) with fish gelatine and ethanol | Sensitive method but performance data not included in manual. Performance data for processed food would be beneficial to form an assessment.
Wheat/Gluten | Sharma, G. M. (2012). "Immunoreactivity and Detection of Wheat Proteins by Commercial ELISA Kits." Journal of Aoac International 95(2): 364-371. (Wheat protein levels at 126% of spiked level) Saito, E., et al. (2019). "The Validation of the Wheat Gluten ELISA Kit." Journal of Aoac International 102(4): 1162-1173. Amnuaycheewa, P., et al. (2022). "Challenges in Gluten Analysis: A Comparison of Four Commercial Sandwich ELISA Kits." Foods 11(5). | Morinaga Wheat/Gluten (Gliadin) ELISA kit II M2114 | Gliadin (Whole wheat protein), polyclonal antibody | Yes | ELISA | Quantitative | ppm (µg protein/ g of food) | Conversion of gliadin to gluten. | 0.31 ppm | 0.31-20 ppm (µg protein/ g of food) | Not stated in manual | Comprehensive range of matrices tested, processed and unprocessed. | (some variability between batches) Millets (Japaneses barnyard millet, foxtail millet, common millet) 0.52-0.84%, rye (>20%), barley(>20%), malt(>20%), oat(>20%), cumin (4.74%), fennel (0.51%), sage (18.8%), red pepper(0.32%), curry powder spice mix 1.74%, thyme (1.59%), dried tangerine peel (0.64%) | None identified | No, although it is permitted to either do a quick or overnight extraction but this is not sample dependent just convenience | Not detailed in manual | Tested using The AOAC Research Institute Performance Tested MethodSM (PTM) | Not identified | Not specified in manual | Sharma et al 2012. Recovert 100-200%, reacted strongly with gliadin fraction rather than glutenin fraction. Amnuaycheewa et al 2022 Responded similarly to the other R5 and G12 methods in the study. | High levels of cross-reactivity with sage. Lower cross reactivity with other herbs. | Method applied to processed and unprocessed foods.
Gluten | - | Rbiopharm RIDASCREEN Total Gluten R7041 | Four different monoclonal antibodies (R5, LMW 1 & 2, HMW) | Yes | ELISA | Quantitative | mg gluten/kg | Yes, glidain to gluten | 4 mg/kg (ppm) gluten depending on matrix | 5 mg/kg (ppm) gluten | Not detailed in manual | Must be non-hydrolysed gluten containing cereals (wheat, rye and barley) in oat products. Storngly acidic or alkaline samples must be neutralised prior to analysis | Manual states .no known cross-reactivities' | Data too sparce to form a judgement. | All samples must be ground to homogenise | AOAC-OMA 2018-15, First Action Status | Interpretation is that, according to manual, AOAC-derived reference material has been used for validation. | None identified | Not detailed in manual | Not detailed in manual | Does not detect D-hordeins. In a composed / processed food (e.g. maize bread) cross reactivities might be different. Interfering substances (e.g. polyphenols) can be detected by spike experiments. Method not applicable to hydrolysed products. | Sensitive method for unhydrolysed foods. Data too sparce to evaluate applicability.
Gluten | - | Biocheck Gluten R5 Allergen-Check ELISA R6098/R6099 | Wheat gliadins, rye secalins and barley hordeins | Yes | ELISA | Quantitative | mg gluten/kg | Not stated in manual | <0.2 mg gluten/kg | 2.5-50 mg gluten/kg | Not detailed in manual | Bio-Check has investigated gluten recoveries form a range of “incurred” samples | Over 65 foods tested, none cross-reacted | Declared data too sparce to evaluate, although 65 (unnamed) matrices tested | No | Not detailed. Reactivity calibrated agains PWG Reference material and FAPAS samples. | Not detailed in manual | Not detailed in manual | Not detailed in manual. States recovery may be diminished due to processing. | Not detailed in manual | List of tested matrices not available. Validation method used not disclosed. | Sensitive, specific, although (long) list of tested matrices not available
Gluten | Sharma, G. M., et al. (2013). "Development of an Incurred Cornbread Model for Gluten Detection by Immunoassays." Journal of Agricultural and Food Chemistry 61(49): 12146-12154. (Recoveries (113−190%)) Amnuaycheewa, P., et al. (2022). "Challenges in Gluten Analysis: A Comparison of Four Commercial Sandwich ELISA Kits." Foods 11(5). | Romer AgraQuant® Gluten G12 10001994 | G12 antibody | Yes | ELISA | Quantitative | ppm gluten | Yes, gluten to gliadin | 2 ppm | 4 ppm | Not detailed in manual | Raw to processed foods, from environmental swabs to rinse water | Not listed | Data too sparce to form a judgement. | For solid samples: homogenize in mortar or blender, add extraction buffer and suspend whereas for liquid samples extraction buffer is added directly. | AACC international method 38-52.01, AOAC official method analysis (OMA) 2014.03 | The AgraQuant® Gluten G12® method used | None identified | Not detailed in manual | Sharma et al 2013. Good precision. Recovery 113-190%. Amnuaycheewa et al 2022 Responded similarly to the other R5 and G12 methods in the study. | Data in manual too sparce to evaluate applicabilty and cross-reactivity. Recovery drops with processing. | AACC international method 38-52.01, AOAC official method analysis (OMA) 2014.03
Gluten | Sharma, G. M. (2012). "Immunoreactivity and Detection of Wheat Proteins by Commercial ELISA Kits." Journal of Aoac International 95(2): 364-371. (Wheat protein levels at 294% of spiked level) | Romer AgraQuant® Gluten 10002002 | Gliadin | Yes | Commercially available | Quantitative | ppm | Yes, gluten to gliadin | 0.6 ppm | 4 ppm | Not detailed in manual | Raw to processed foods, from environmental swabs to rinse water | Barley 5%, oat 0.028%, rye 100%, triticale 4% and wheat 100% | Data too sparce to form a judgement. | For solid samples: homogenize in mortar or blender, add extraction buffer and suspend whereas for liquid samples extraction buffer is added directly. For samples outside of 6-8 pH, pH should be adjusted with NaOH or HCl. For samples with high tannins or polphenols add 10% fish gelatin |  | Approved as AOAC Official Method | Yes, compared to RBiopharm R7051 kit to determine gluten in wheat. Results anonymised. Both kits showed approx. 100% recovery and resulted varied significantly depending on wheat cultivar. | Not detailed in manual | Sharma et al 2012. Recovery too high, met LOQ stipulated in manual.Reacted strongly with gliadin fraction compared to glutenin fraction. Sharma et al 2013. Recovery 113-190%. Good precision. Recovery reduced with processing. Despite using a different anttibody, responded similarly to a G5 kit in a study to determine wheat cultivars with 100% recoveries for certain cultivars. | Recovery (in lab evaluation) reduced with processing). Data in manual too sparce to evaluate applicabilty and cross-reactivity. | Sensitve method but performnace data too sparce to evaluate.
Gluten | Gomaa, A. and J. I. Boye (2013). "Impact of thermal processing time and cookie size on the detection of casein, egg, gluten and soy allergens in food." Food Research International 52(1): 483-489. Amnuaycheewa, P., et al. (2022). "Challenges in Gluten Analysis: A Comparison of Four Commercial Sandwich ELISA Kits." Foods 11(5). | Neogen Veratox for Gliadin R5 8510 | Gliadin R5 | Yes | ELISA | Quantitative | ppm gliadin | Yes, gliadin to gluten | From AOAC validation: Cereal–2.16 ppm gliadin; cooked hamburger–1.21 ppm gliadin; bread–0.63 ppm gliadin; and rice flour–1.23 ppm gliadin. Neogen internally validated commodity LODs were all under 2.5 ppm gliadin (defined as the mean value of 10 negative samples + 3.3 standard deviations) | From AOAC validation:Cereal–5.71; cooked hamburger–3.15; bread–1.84; and rice flour–3.31. Neogen internally validated commodity LOQs were all under 2.5 ppm (described as the lowest concentration point on the calibration curve that this test can reliably detect gliadin). | Not detailed in manual | Rice flour, cereal, bread and cooked hamburger were validated by AOAC Research Institute | None stated | Data too sparce to form a judgement. | Dependent on whether sample is heat processed, if it is an unknown commodity | AOAC | Analysis (OMA) 2014.03 and as AACC | Not detailed in manual | Not detailed in manual | Amnuaycheewa et al 2022 Responded similarly to the other R5 and G12 methods in the study, althoughreported differently for yeast extract, hemp protein powder and cookie. | None identified | Sensitive method, validated by AOAC
Gluten | - | Eurofins Sensispec SENSISpec INgezim Gluten R5 | R5 Monoclonal Antibody, which is specific for proteins from wheat, rye and barley. | Yes | ELISA | Quantitative | ppm gluten | Yes, gluten to gliadin | 3 ppm | 25 ppm | Not detailed in manual | Not detailed in manual | Not detailed in manual | Data too sparce to form a judgement. | No | AOAC | International Approved Method 38 | Yes | Recovery data and precision data not included in online-instructions. | None identified | None identified | Sensitive method, validated by AOAC
Gluten | - | RBioPharm SureFood® ALLERGEN Gluten Art. No.: S3606 | DNA of gluten-containing cereals (wheat such as spelt and khorasan wheat, rye, barley, oats) according to directive (EC) 1169/2011 | No | PCR | Quantitative | mg/kg | mg allergenic substance /kg food (NB not mg allergenic protein/kg food) | ≤ 0.4 mg / kg | 1 mg / kg using SureFood® PREP Advanced, protocol 1 | Corn flour | Not detailed in manual | Not detailed in manual | Data too sparce to form a judgement. | Advanced Protocol 1 recommended for this, other matrices may require different methods (see paperwork) | Not stated in inset | No | The technical validation of instruments was performed on Roche LightCycler® 2.0, Roche LightCycler® 480 II, Roche cobas® z 480 Analyzer, Qiagen Rotor-Gene Q, Applied Biosystems 7500, Bio-Rad CFX96, Bio Molecular Systems MIC, LTF MyGo Pro, Agilent AriaDx and Agilent Mx3005P | Recovery data and precision data not included in online-instructions. | None identified | Does not report in mg allergen per kg food. Performance data, including applicability, too sparce to evaluate. | Does not measure mg allergenic protein as is the preferred reporting format.
Gluten | - | ElisaSystems Generation 3 Gluten Residue Detection Kit. Ref. ESGLT3 | Gliadin and Gluten protein | Yes | ELISA, based on 2D4 antibody 'with R5-like specificity' | Quantitative and qualitative | mg/kg | Multiply by 2 to transfer from gliadin to gluten | 1ppm gliadin, 2 ppm gluten mg/kg | 2 ppm gliadiin, 4ppm gluten | None detailed in manufacturer's user guide | Raw and cooked foods, nutritional supplements, and environmental samples for the presence of both native and deamidated gluten residues derived from wheat, barley, rye, and their hybrid. Nature of matrices not detailed. Not suitable for analysis of oat. | None detailed in manufacturer's user guide | Specific matrices not listed. | No | This ELISA Systems product was manufactured in a facility that maintains a quality management system certified to ISO9001:2015 | Not detailed in manufacturer's user guide. | Not detailed in manufacturer's user guide. | Not detailed in manufacturer's user guide. | Not detailed in manufacturer's user guide. | Nature of matrices not detailed. User guide states that kit is designed for screening only and not for confirmation. | Medium length ELISA method, no information about matrices tested or validation data or false positive data. Not suitable for analysis of oat. The ELISA SYSTEMS Oat Protein kit is more suitable for products where the specific detection of oat residues is required. User guide states that kit is designed for screening only and not for confirmation.

## Crustacea
Allergen type | Citing Papers | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target e.g. protein/ peptide/ DNA target. Food or sub-food e.g. whole egg or egg white protein only. Name the target if this information is available e.g. put protein name, gene name or sequence info if available. | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredient | List relevant/pertinent untested matrices | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data (inc. precision, recovery). This will comne both from literautre review (if available and also from PTG) | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Crustacean | - | ELISAKits Crustacean (tropomyosin) ELISA kit | Tropomyosin protein | Yes, measures allergenic protein. | ELISA | Quantitative | ppm | Not stated in manual | 0.09 ppm | 20.0 ppm | Not detailed on web site. | Fish, meat, baked goods and soups, soy sauce. | The following displayed some cross reactions: Cockroach (protein) (0.01%). No cross reaction was detected for the following: Milk, Fish, Macadamia nut, Egg, Oyster, Chestnut, Wheat, Sunflower seeds, Pine nut, Rye, Pumpkin seeds, Soy, Oats, Cashew, Lecithin (soy), Barley, Peanut, Pea, Rice, Hazelnut, Bean, Corn, Almond, Potato, Buckwheat, Pecan, Carrot, Sesame, Coconut, Leek, Pork meat, Brazil nut, Celery, Beef, Walnut, Mustard, Chicken meat, Pistachio. | Sauces, soup. | None identified. None detailed on website. | None detailed on website. | None detailed on website. | Intra-Assay Precision: 6 – 8%, inter-assay precision 5-12%. Recovery: Meat (97%), Fish (93%), Vegetable soup (93%), Bakery products (90%), Soy sauce (84%). Linearity: 74 – 114% | None detailed in manual | May cross-react with insect protein | Peer-reviewed manuscripts detailed regarding method development, precision data included which is rare for commercial kits.
Crustacean | - | Eurofins SENSISpec ELISA Crustaceans (Tropomyosin) (HU0030006 / HU0030030) | Tropomyosin protein from Penaeus indicus | Yes, measures allergenic protein. | ELISA | Quantitative | ppb | Not stated in manual | 0.9-10.3 ppb, dependant on matrix. | 20 ppb | soy sauce, vegetable soup, bakery goods, fish, meat spiked for LOD, LOQ and recovery studies | Instruction manual lists fish products, soups, dressings, bakery products and meat products | Cockroach. 38 ingredients tested with no cross-reactivity. | There are considerations regarding whether novel food insects will cross-react e.g. black soldier fly | Not referenced in instruction manual | Not applicable | Not referenced in instruction manual | Intra-assay precision 6-8%, inter-assay precision 5-12%. Recovery: 84-97%, depending on matrix. | Not referenced in instruction manual | There will be future considerations regarding cross-reactivity with novel food insects if these are authorised as food in UK. | Conversion factors are available for 6 species of seafood, both cooked and raw, so species present needs to be known for an accurate quantitaive .
Crustacean | Jeong, S. G. and S. H. Kim (2020). "Application of commercial kits using DNA-based and immunochemical methods for determination of shrimp allergens in kimchi and its ingredients." Journal of Food Science 85(10): 3638-3643. | Rbiopharm RIDASCREEN FAST crustacean R7312 | Antibody specifically detects crustacean-proteins (mainly tropomyosin) | Yes, measures allergenic protein. | ELISA | Quantitative | mg/kg | Not stated in manual | 2 mg/kg | 20 mg/kg | Not disoclosed in manual | In a composed / processed food (e.g. corn bread) cross reactivities might be different. Interfering substances (e.g. polyphenols) can be detected by spike experiments | Mustard, curcuma, beans, mussels and arthropods. 'In a composed / processed food (e.g. corn bread) cross reactivities might be different. Interfering substances (e.g. polyphenols) can be detected by spike experiments.' | None identified | Not referenced in instruction manual | Not referenced in instruction manual | Not referenced in instruction manual | Not referenced in instruction manual | Jeong and Kim 2020 Kit capable of detecting allergen at eliciting levels. | There will be future considerations regarding cross-reactivity with novel food insects if these are authorised as food in UK. | Indication is that a wide range of matrices have been tested (due to unusual matrices listed as cross-reacting. Further validationinformation required.
Crustacean | - | Biocheck R6027/R6026 | Polyclonal antibodies to tropomyosin Pen a1 | Yes, measures allergenic protein. | ELISA | Quantitative | µg tropomyosin/kg | Yes, from µg tropomyosin/kg to mg crustacean tissue/kg, on a dry weight basis | 0.1 µg tropomyosin/kg | 20 µg tropomyosin/kg | Not disoclosed in manual | Included incurred processed matrices | Cockroach protein (0.01%) | None identified | Not referenced in instruction manual | Not referenced in instruction manual | Not referenced in instruction manual | Recovery 84-97% tested in 5 matrices | Not referenced in instruction manual | No reference to AOAC guideline use during development,no inter-lab trial data. | Fairly detailed validation data presented, good specificity and sensitivity, applicable to processed incurred products and reports data in terms of level of allergen per kg food.
Crustacean | - | Morinaga Crustacean ELISA kit II M2118 | Black Tiger Tropomyosin (Whole crustacean protein) | Not detailed in maunal | ELISA | Quantitative | ppm | Not stated in manual | 0.31 ppm | 0.31-20 ppm (ug protein/ g of food) | Not disoclosed in manual | Processed and raw tested. | Squid. May show cross-reaction with fish, shellfish, seaweed due to predation of shrimp and crub. | None identified | Not referenced in instruction manual | Not referenced in instruction manual | Not referenced in instruction manual | Not referenced in instruction manual | Not referenced in instruction manual | There will be future considerations regarding cross-reactivity with novel food insects if these are authorised as food in UK. | Testing of method seems comprehensive fro available data.
Crustacean | Jeong, S. G. and S. H. Kim (2020). "Application of commercial kits using DNA-based and immunochemical methods for determination of shrimp allergens in kimchi and its ingredients." Journal of Food Science 85(10): 3638-3643. | Romer AgraQuant® Crustacea 10002076 | Crustacean tropomyosin | Not detailed in maunal | ELISA | Quantitative | ppb | Not stated in manual | 0.9 ppb | 20 ppb | Not disoclosed in manual | raw to processed foods, from environmental swabs to rinse water | Cockroach (protein) 0.01% | Data too sparce to evaluate | Not referenced in instruction manual | Not referenced in instruction manual | Not referenced in instruction manual | Not referenced in instruction manual | Jeong and Kim 2020 Kit capable of detecting allergen at eliciting levels. | There will be future considerations regarding cross-reactivity with novel food insects if these are authorised as food in UK. | Little validation data available to fully evaluate.
Crustacean | Jeong, S. G. and S. H. Kim (2020). "Application of commercial kits using DNA-based and immunochemical methods for determination of shrimp allergens in kimchi and its ingredients." Journal of Food Science 85(10): 3638-3643. | Neogen Veratox for crustacea allergen 8520 | Crustacea protein residue | Not detailed in maunal | ELISA | Quantitative | ppm | Not stated in manual | 2.5 ppm | 2.5 ppm | Not disoclosed in manual | Not stated in manual | No | Data too sparce to evaluate | Not referenced in instruction manual | Not referenced in instruction manual | Not referenced in instruction manual | Intra-assay variabliliy 3.0%, inter-assay variability 4.1% recovery mean 83% | Jeong and Kim 2020 Kit capable of detecting allergen at eliciting levels. | There will be future considerations regarding cross-reactivity with novel food insects if these are authorised as food in UK. | Little validation data available to fully evaluate.
Crustacean | Jeong, S. G. and S. H. Kim (2020). "Application of commercial kits using DNA-based and immunochemical methods for determination of shrimp allergens in kimchi and its ingredients." Journal of Food Science 85(10): 3638-3643. | RBioPharm SureFood Allergen Crustaceans S3612 | Crustacea DNA | Not detailed in maunal | PCR | Quantitative | mg/kg | Not stated in manual | ≤ 0.4 mg / kg | 1 mg / kg using SureFood® PREP Advanced, protocol 1 | Not disoclosed in manual | The limit of detection and quantification of the complete method (DNA extraction and real-time PCR) depends on sample matrix, processing grade, DNA-preparation and DNA-content - different extraction protocols required for different matrices | Cross reactivity was observed with DNA extracts from abalone (Haliotis spp.) with 100 % | Data too sparce to evaluate | Not referenced in instruction manual | Not referenced in instruction manual | The technical validation of instruments was performed on Roche LightCycler® 480 II, Roche cobas® z 480 Analyzer, Qiagen Rotor-Gene Q, Applied Biosystems 7500, Bio-Rad CFX96, Bio Molecular Systems MIC, Agilent AriaDx and Agilent Mx3005P | Data too sparce to evaluate | Exact part number not specified by Jeong and Kim 2020 founda crustacea kit by this manufacturer kit to have only qualitative capability. | Validation data available on request | Data too sparce to evaulate method suitability on a range of matrices.
Crustacean | - | RBioPharm SureFood Allergen Crustaceans S3613 | Crustacea DNA | Not detailed in maunal | PCR | Quantitative | mg/kg | Not stated in manual | ≤ 0.4 mg / kg | Not detailed in manual | Not disoclosed in manual | For DNA preparation, the use of SureFood® PREP Advanced (S1053), protocol 1, is recommended | Not detailed in manual | Data too sparce to evaluate | Not referenced in instruction manual | Not referenced in instruction manual | Internal validation was performed on Roche LightCycler® 2.0, Roche LightCycler® 480 II, Roche cobas® z 480 Analyzer, Qiagen Rotor-Gene Q, Applied Biosystems 7500, Bio-Rad CFX96, Bio Molecular Systems MIC, AriaDx and Agilent Mx3005P | Data too sparce to evaluate | Exact part number not specified by Jeong and Kim 2020 founda crustacea kit by this manufacturer kit to have only qualitative capability. | Validation data available on request | Data too sparce to evaulate method suitability on a range of matrices.

## Egg
Allergen type | Citing Papers | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target e.g. protein/ peptide/ DNA target. Food or sub-food e.g. whole egg or egg white protein only. Name the target if this information is available e.g. put protein name, gene name or sequence info if available. | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredients | List relevant/pertinent untested matrices | Is there more than one extraction method dependant on matrix | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data. | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Egg | - | ELISAKits ovalbumin ELISA kit | Ovalbumin egg white protein | Total ovalbumin (one of the egg allergenic proteins) measured | ELISA | Quantitative | ppb | Not stated in manual | 4 ppb | 25 ppb | Not detailed on website | Wine | Specificity (Cross Reactivity): Egg White Protein (total) (75%), Conalbumin (< 0.2%), Lysozyme, Ovomucoid (< 0.02%), Bovine Gelatine, Fish, Isinglass, Non Fat Dry Milk, Fish Gelatine (0%). | Kit specific for wine. Other foods not tested. | Not detailed on website | Assume not, none detailed on website | Assume not, none detailed on website | None identified. | Intra-Assay Precision: 3 – 4%, Inter-Assay Precision: 2 – 3%, Linearity 94-96%. Recovery After spiking wine samples with different amounts of ovalbumin, 95-115% recovery results were obtained. | None identified. | Applies only to wine. | Peer-reviewed manuscripts detailed regarding method development, precision data included which is rare for commercial kits.
Egg | - | SENSISpec ELISA Lysozyme (HU0030012 / HU0030036) | Antibody directed against lysozyme' protein | Only a method to detect lysozyme is provided in the instruction manual | ELISA | Quantitative | ppb | Only a method to detect lysozyme is provided in the instruction manual | 2 ppb for wine, 10 ppb for other matrices. | 25 ppb for wine, 125 ppb for other matrices. | Wine, cheese. | Wine, cheese. | Egg white total protein 2.2% cross-reactivity. Very low cross-reactivity with cashew, conalbumin, chicken, ovalbumin, ovomucoid. 69 additional ingredients tested with no cross-reactivity. | N/A | Yes, method for wine and method for other matrices. | None detailed in instruction manual | None detailed in instruction manual | None detailed in instruction manual | Intra-assay precision 2-6%, inter-assay precision 2-3%. Recovery 85-99%, determined in wine and cheese. | None listed in instruction manual. | Only specific to wine and cheese matrices | Specific to wine and cheese. Cross-reactivity studies are comprehensive
Egg | Parker, C. H., et al. (2015). "Multi-allergen Quantitation and the Impact of Thermal Treatment in Industry-Processed Baked Goods by ELISA and Liquid Chromatography-Tandem Mass Spectrometry." Journal of Agricultural and Food Chemistry 63(49): 10669-10680. | Rbiopharm Ridascreen fast Egg protein R6402 | Egg white proteins ovalbumin and ovomucoid | Egg white protein | ELISA | Quantitative | mg/kg | can convert from whole egg powder to total protain | 0.15 mg egg white/kg food | 0.4 mg egg white protein/kg food | Whole egg powder NIST reference material 8445 ice cream, noodles, salad dressing, sausages, baking mixtures for cakes or bread and wine. It can be assumed that the assay is also suitable for the analysis of other foods, however, this must be verified by the user before applying the test kit to that food category | No cross-reactivity with raw or cooked meat of chicken, turkey, hen, pork and beef. | No cross reactivity was observed to raw or cooked meat of chicken, turkey hen, pork and beef. | Information too sparce to evaluate.Not detailed in manual | No | None detailed in instruction manual | None detailed in instruction manual | Collaborative study (18 different laboratories) for egg white in wine resulted in recovery between 88 – 100.5 % (red wine) and 70.4 – 86 % (white wine) from manufacture's website | Recovery in pasta, biscuit, cookies, dark chocolate and sausage ranges from 82-98% byt not information regarding whether these were raw, spiked or incurred. | None listed in instruction manual. | In a composed / processed food (e.g. corn bread) cross reactivities might be different. Interfering substances (e.g. polyphenols) can be detected by spike experiments. For heat-treated samples, recommend use of manufacturers' alternative kit (Ridascreen Egg) | More data required regarding relevent matrices
Lysozyme | - | Rbiopharm RIDASCREEN®FAST Lysozym - 6452 | Lysozyme | Egg white protein | ELISA | Quantitative | mg/kg | Not stated in manual | 0.15 mg egg white/kg food | 0.4 mg egg white protein/kg food | Wine (more sensitive in this matrix), cheese and sausage | Tested on 'pure' raw ingredients which are not listed in the manual. | Only one food tested for cross-reactivity mentioned in manual. Need to apply for validation report. | Information too sparce to evaluate.Not detailed in manual | Different methods wine, and cheese and sausage | Not detailed in manual | Not detailed in manual | Not detailed in manual | Recovery in pasta, biscuit, cookies, dark chocolate and sausage ranges from 82-98% byt not information regarding whether these were raw, spiked or incurred. | Not detailed in manual | Cross reactivities of the used antibodies have been determined for the pure food (e.g. corn flour). In a composed / processed food (e.g. corn bread) cross reactivities might be different. Interfering substances (e.g. polyphenols) can be detected by spike experiments. | Applies to wine, cheese and sausage only.
Egg | Heick, J., et al. (2011). "Application of a Liquid Chromatography Tandem Mass Spectrometry Method for the Simultaneous Detection of Seven Allergenic Foods in Flour and Bread and Comparison of the Method with Commercially Available ELISA Test Kits." Journal of Aoac International 94(4): 1060-1068. | Rbiopharm Ridascreen fast Egg R6411 | Egg protein | Egg white protein | ELISA | Quantitative | mg/kg | Manual states can convert from whole egg powder to total protein and to egg white protein. This would rely on the food matrix containing whole egg. | 0.15 mg egg white/kg food | 0.4 mg egg white protein/kg food | Not stated in manual | Ice cream, noodles, salad dressing, wine, chocolate, and cookies. | Egg from gooses, quails and ostrich | Suitability of test on foods not listed will need to be checked by the user | Two different extraction methods in one kit for the detection of processed egg with A-AEP or the easier extraction for native egg with the AEP. | AOAC-RI validation (PTM) in preparation | AOAC-RI validation (PTM) in preparation | AOAC-RI validation (PTM) in preparation | Recovery in pasta, biscuit, cookies, dark chocolate and sausage ranges from 82-98% but not information regarding whether these were raw, spiked or incurred. | Not detailed in manual | The cross-reactivities of the antibodies used were determined for pure food (e.g. cornmeal). These cross-reactivities may be different in compound / processed foods (e.g. cornbread). Spiking experiments can be used to detect potentially interfering substances (e.g. polyphenols) | Sensitive and specific
Egg | - | Biocheck R6033, R6032 | Ovomucoid | Egg white protein | ELISA | Quantitative | mg/kg egg white protein | Not stated in manual | 0.15 mg egg white/kg food | 0.4 mg egg white protein/kg food | Not stated in manual | 36 matrices tested | There is a low level of cross-reactivity with raw chicken meat but not when cooked (95°C). Egg proteins conalbumin (2.6%) and ovalbumin (0.25%) | None, although not clear wheather matrices were cooked, raw, spiked or incurred | Under development, although on line presence. | Not detailed in manual | Not detailed in manual | Not detailed in manual | Recovery in pasta, biscuit, cookies, dark chocolate and sausage ranges from 82-98% byt not information regarding whether these were raw, spiked or incurred. | Not detailed in manual | Cross-reacts with chicken, specific to egg white proteins, not yolk. | Sensitive and specific, tested on a wide range of matrices
Egg | - | Romer AgraQuant® Ovalbumin 10002092 | Highly purified polyclonal antibodies raised against ovalbumin | Yes | ELISA | Quantitative | ppb | Not stated in manual | 4 ppb | 25 ppb | Not stated in manual | Raw to processed foods, from environmental swabs to rinse water | Conalbumin <0.2%, chilli 0.000004%, cumin 0.000004%, lysozyme <0.02%, ovomucoid <0.02% and whole egg powder 15% | Details in manual too sparce to evaluate | For solid samples: homogenize in mortar or blender, add extraction buffer and suspend whereas for liquid samples extraction buffer is added directly. For samples outside of 6-8 pH, pH should be adjusted with NaOH or HCl. For samples with high tannins or polphenols add 10% fish gelatin | Not detailed in instructions | N/A | Not detailed in manual | Not detailed in manual | Not detailed in manual | Specific to egg white | More validation information required to form an opinion.
Egg | - | Romer AgraQuant® Lysozyme 10002090 | Highly purified polyclonal antibodies raised against lysozyme | Yes | ELISA | Quantitative | ppb | Not stated in manual | 2 ppb | 25 ppb | Not stated in manual | Raw to processed foods, from environmental swabs to rinse water | Conalbumin 0.0001%, total egg white protein 2.2%, ovalbumin 0.0001% and ovomucoid 0.0001% | Details in manual too sparce to evaluate | For solid samples: homogenize in mortar or blender, add extraction buffer and suspend whereas for liquid samples extraction buffer is added directly. For samples outside of 6-8 pH, pH should be adjusted with NaOH or HCl. For samples with high tannins or polphenols add 10% fish gelatin | Not detailed in instructions | N/A | Not detailed in manual | Not detailed in manual | Not detailed in manual | Specific to egg white | More validation information required to form an opinion.
Egg | - | Romer AGRAQUANT ALLERGEN - EGG WHITE 10002026 | Egg white protein | No | ELISA | Quantitative | ppm | Not stated in manual | 0.05 ppm | 0.4 ppm | Not stated in manual | Raw to processed foods, from environmental swabs to rinse water | Raw chicken meat <0.001% | Information too sparce to assess | For solid samples: homogenize in mortar or blender, add extraction buffer and suspend whereas for liquid samples extraction buffer is added directly. For samples outside of 6-8 pH, pH should be adjusted with NaOH or HCl. For samples with high tannins or polphenols add 10% fish gelatin | Not detailed in manual | Not detailed in manual | Not detailed in manual | Not detailed in manual | Not detailed in manual | Insufficient detail provided to evaluate | Insufficient information provided to evaluate
Egg | Fu, T. J., et al. (2010). "Effect of Heat Treatment on the Quantitative Detection of Egg Protein Residues by Commercial Enzyme-Linked Immunosorbent Assay Test Kits." Journal of Agricultural and Food Chemistry 58(8): 4831-4838., Gomaa, A. and J. I. Boye (2013). "Impact of thermal processing time and cookie size on the detection of casein, egg, gluten and soy allergens in food." Food Research International 52(1): 483-489., Heick, J., et al. (2011). "Application of a Liquid Chromatography Tandem Mass Spectrometry Method for the Simultaneous Detection of Seven Allergenic Foods in Flour and Bread and Comparison of the Method with Commercially Available ELISA Test Kits." Journal of Aoac International 94(4): 1060-1068., Parker, C. H., et al. (2015). "Multi-allergen Quantitation and the Impact of Thermal Treatment in Industry-Processed Baked Goods by ELISA and Liquid Chromatography-Tandem Mass Spectrometry." Journal of Agricultural and Food Chemistry 63(49): 10669-10680. | Morinaga Egg (Ovalbumin) ELISA Kit II M2111 | Ovalbumin (for whole egg protein) | Yes | ELISA | Quantitative | ppm | Not stated in manual | 0.31 ppm | 0.31-20 ppm (ug protein/ g of food) | Not stated in manual | Not stated in manual | None dectected - see documentation for full list checked, >50 matrices tested | None identified | No, although it is permitted to either do a quick or overnight extraction but this is not sample dependent just convenience | Not detailed in manual | Not detailed in manual | Not detailed in manual | Not detailed in manual | Fu, 2010 comparison to biokit and veratox for thermally processed eggs, responsed best to thermal processing compared to two other kits. Gomaa and Boye 2013 this kit responds better than other kit tested for recovery following thermal processing. | Insufficient detail provided to evaluate | Specific to a wide range of tested matrices. Seems to perform best compared to other kits for thermally processed products (according to the literature).
Egg | - | SENSISpec ELISA Egg White 96/48 SENSISpec ELISA Egg White 96/48 Tests (Cat.nr. HU0030007 / HU0030031) | Egg white (ovomucoid), hen's egg | Yes, to whole egg powder | ELISA | Quantitative | ppm | Not stated in manual | 50 ppb | 0.4 ppm egg white | Pasta, bakery products, sausage and chocolate | Pasta, bakery products, sausage and chocolate | None listed, 53 ingredients tested. NO significant cross-reactions with chicken or duck or with other egg proteins. | None identified | No | Not detailed in manual | Not detailed in manual | Not detailed in manual | Intra-assay precision 4-9%, inter-assay precision 3-7%. Recovery 82-98% for selected matrices. | Not detailed in manual | Insufficient information provided to evaluate | Target protein (ovomucoid) is heat stable and can resist common production processes like baking. Not applicable to egg yolk.
Egg | Williams, K. M., et al. (2004). "Determination of egg proteins in snack food and noodles." Journal of Aoac International 87(6): 1485-1491., Fu, T. J., et al. (2010). "Effect of Heat Treatment on the Quantitative Detection of Egg Protein Residues by Commercial Enzyme-Linked Immunosorbent Assay Test Kits." Journal of Agricultural and Food Chemistry 58(8): 4831-4838., Gomaa, A. and J. I. Boye (2013). "Impact of thermal processing time and cookie size on the detection of casein, egg, gluten and soy allergens in food." Food Research International 52(1): 483-489., Parker, C. H., et al. (2015). "Multi-allergen Quantitation and the Impact of Thermal Treatment in Industry-Processed Baked Goods by ELISA and Liquid Chromatography-Tandem Mass Spectrometry." Journal of Agricultural and Food Chemistry 63(49): 10669-10680. | Neogen - Veratox for egg allergen 8450 | Egg white protein | Yes dried whole egg | ELISA | Quantitative | ppm | Not stated in manual | 2.5 ppm whole dried egg | 2.5 ppm whole dried egg | Whole dried egg | Pasta, salad dressing, cake mix and ice cream | No | Information too sparce to assess | No | Not detailed in manual | Not detailed in manual | Not detailed in manual | Intra-assay precision 2.8%, inter-assay precision3.6%. Recovery 88% | Fu, 2010 10.1021/jf903536z comparison to biokit and morinaga for thermally processed eggs, thermal processing reduced recovery. Williams et al 2004 processing has a significant effect on detection levels. Gomaa and Boye 2013 this kit responded less favourably than other kit tested for recovery following thermal processing. | Concentrated food additives,colors and flavors may cause interferences on ELISA test methods. Hydrolyzed and fermented proteins may not be detected using ELISA methods for allergen testing. | Insufficient information provided to evaluate

## Fish
Allergen type | Citing Papers | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target e.g. protein/ peptide/ DNA target. Food or sub-food e.g. whole egg or egg white protein only. Name the target if this information is available e.g. put protein name, gene name or sequence info if available. | Is method commercially available, or does it have sufficient validation data? | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredient | List relevant/pertinent untested matrices | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data. False reporting rate and measurement uncertainty data (inc. precision, recovery). | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Fish | - | ELISAKits (Parvalbumin) ELISA Kit | Parvalbumin protein | Commercially available | Yes | ELISA | Quantitative | ppm (cod) | Can't interconvert between species | 1.4 ppm | 4.0 ppm | Details not available online. | Manual states: 'sauces, soups, crackers, wine, Asian food and surimi'. No cross reaction was detected for the following: Almond, Hazelnut, Pork, Barley, Isinglass, Potato, Bean, Lamb, Pumpkin seed, Beef, Macadamia, Rice, Brazil nut, Milk, Rye, Buckwheat, Millet, Scallop, Carrot, Mustard, Sesame, Cashew, Oat, Shrimp, Celery, Onion, Soy, Chicken, Pea, Sunflower seed, Corn, Peanut, Walnut, Egg, Pecan, Wheat, Fish gelatin, Pistachio. | Details not available online. | Information too sparce to form an opinion. | Details not available online. | Details not available online. | Not detailed in manual | Intra-Assay Precision: 7 – 12%, inter-assay precision: 4-10%. Linearity: 89-105%. Recovery: Soup (117%), Surimi (114%), Worcester Sauce (112%), Wine (red)(103%), Asia Sauce (103%), Cracker (99%), Spring Roll (93%). | Details not available online | Information too sparce to form an opinion. Recovery will likley be dependant on species. | Tested against a relevent range of foods. Outcome of test relies on prior knowledge of species.
Fish | - | SENSISpec ELISA Fish (Parvalbumin) (HU0030008 / HU0030032) | Parvalbumin protein | Commercially available | Not detailed in manual | ELISA | Quantitative | ppm (cod) | Conversion factors available for 27 fish species (fresh fish). For cooked fish, the ELISA response was 25% leass when tested on cod. | 1.4 ppm cod (0.3-1.8 ppm cod depending on matrix) | 4 ppm cod | Not detailed in manual | Wine, soups, sauces, crackers, surimi and similar products' are listed in instruction manual | 38 ingredients tested with no cross-reactivity | None identified, wine, soup, sauces, baked goods, surimi and spring roll evaulated. | None referenced in instruction manual | Details not available online. | Not detailed in manual | Intra-assay precision is 7-12%, inter-assay precision is 4-10. Recovery is 93-117% depending on matrix. | Not detailed in manual | Information too sparce to form an opinion. Recovery will likley be dependant on species. | Weakness: conversion data available for 27 species of raw fish and for one sppecies of cooked fish. Publications available in instructions.
Fish | - | Rbiopharm SureFood Allergen Fish S3610 | Fish DNA according to (EC) 1169/2011 (the group of fishes contains about 30000 species) | Commercially available | Not detailed in manual | PCR | Quantitative | mg/kg | Can't interconvert between species | ≤ 1 mg / kg | 4 mg / kg using SureFood® PREP Advanced, protocol 1 | Corn flour | Not detailed in manual | Cross reactivity was observed with DNA extracts from muscovy duck (Cairina moschata). | Information too sparce to form an opinion. | None referenced in instruction manual | Details not available online. | Not detailed in manual | Not detailed in manual | Not detailed in manual | Information too sparce to form an opinion. Recovery will likley be dependant on species. | Information too sparce to form an opinion. Recovery will likley be dependant on species.
Fish | - | Biocheck | Parvalbumin allergen, using antibodies to cod parvalbumin | Commercially available | Kit provides level of cod in a sample. | ELISA | Qualitative and Quantitative | mg/kg | Not detailed in manual. | <0.5 mg/kg cod | 5-125 mg/kg cod | Fresh cod | Range of fresh fish species. | Buckwheat (0.02%), amaranth (0.01%), soya (0.004%). 50 commodities tested (but not listed). | Soup, pie, risotto | Manual states 'validation broadly followied AOAC Food Allergen Community Guidance' | Details not available online. | Not detailed in manual | Kit states that duplicate precision is <7.5% repeatabilty of absorbance and that 'duplicate precision <15% relative standard deviation mg/kg cod' | Not detailed in manual | 100% reactive with cod, 31-57% reactive with other Gadiformes, 7-8% reactive with Salmonifirmes, 0.16% reactive with tuna and 0.01% reactive with monkfish. Thereofre, the species of fish must be known in order to estimate the level of fish present in a sample. Procewssing may significnatly reduce the reactivity of the assay, according to the manual. Kit is 100% reactive to cod when raw, was 8% reactive when boiled in milk. | Require prior knowledge of fish species present to quantify allergen.Much diminished response of the species is not cod.Significnat loss in reactivity with thermal processing.
Fish | Ruethers, T., et al. (2020). "Commercial fishELISAkits have a limited capacity to detect different fish species and their products." Journal of the Science of Food and Agriculture 100(12): 4353-4363. | Romer AGRAQUANT ALLERGEN - FISH 10002083 | Parvalbumin protein | Commercially available | Not detailed in manual | ELISA | Quantitative | ppm (cod) | Not detailed in manual. | 1.4 ppm | 4 ppm | Fresh cod | If processed food is analyzed, the grade of processing needs to be taken into account (e.g. cod meat cooked for 20 minutes resulted in a 25% reactivity compared to fresh cod). | Squid, 0.001% | Information too sparce to form an opinion. | None referenced in instruction manual | Details not available online. | Not detailed in manual | Data too sparce to evaulate | Not detailed in manual | If processed food is analyzed, the grade of processing needs to be taken into account (e.g. cod meat cooked for 20 minutes resulted in a 25% reactivity compared to fresh cod). | Outcome of test relies on prior knowledge of species. Tested matrices unclear.

## Lupin
Allergen type | Citing Papers | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target e.g. protein/ peptide/ DNA target. Food or sub-food e.g. whole egg or egg white protein only. Name the target if this information is available e.g. put protein name, gene name or sequence info if available. | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredients | List relevant/pertinent untested matrice | Is there more than one extraction method dependant on matrix (e.g. sometimes the case for ELISA)? Need to detail all extraction methods and associated data. | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data (inc. precision, recovery). | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Lupin | - | SENSISpec (HU0030011 / HU0030035) | Lupin protein | Not specified in manual | ELISA | Quantitative | ppm | Not specified in manual | 0.2 ppm | 2 ppm | Sausage, bakery products, potato products, ketchup and juices (LOD 0.1-0.7 ppm) | Sausage, bakery products, potato products, ketchup and juices | Lupin protein fractions cross-react with soy and peanut allergens | None | No | Not stated in manual | Not stated in manual | Not stated in manual | Intra-assay precision, inter-assay precision 6-6%, 5-12%, recovery 98-113% | Not stated in manual | Cross-reacts with peanut and soy allergens. Tested on a wide variety of matrices, no information regarding whether these matrices were spiked or incurred. | Sensitive, cross-reacts with peanut and soyabean
Lupin | Röder, M., et al. (2013). "Detectability of Lupine Seeds by ELISA and PCR May Be Strongly Influenced by Potential Differences between Cultivars." Journal of Agricultural and Food Chemistry 61(25): 5936-5945., Koeberl, M., et al. (2018). "Lupine allergen detecting capability and cross-reactivity of related legumes by ELISA." Food Chemistry 256: 105-112. | Rbiopharm RIDASCREEN FAST Lupine R6102 | Antibody specifically detects sweet lupine proteins, including the γ-conglutin as well as all European species of sweet lupines (Lupinus albus, luteus and angustifolius) relevant in the food and feed industry | Not specified in manual | ELISA | Quantitative | mg lupin protein /kg | Not specified in manual | 0.32 mg / kg (ppm) lupine protein (mean) 0.12 – 0.65 mg / kg (ppm) lupine protein* *depending on matrix | 1 mg / kg (ppm) lupine protein | Not specified in manual | Juices, wine, beer, sausages, noodles, chocolate, bakery products or ice cream | Chickpea, soy flour, roasted hazelnut, curcuma, curry and fenugreek | None | No | Not stated in manual | Not stated in manual | Not stated in manual | Not stated in manual | Röder et al 2013, Very high differences in recovery between cultivars. | Specific and sensitive for a wide range of matrices. More data required regarding response to different cultivars. | Specific and sensitive for a wide range of matrices. More data required regarding response to different cultivars.
Lupin | Koeberl, M., et al. (2018). "Lupine allergen detecting capability and cross-reactivity of related legumes by ELISA." Food Chemistry 256: 105-112. | Romer AgraQuant® Lupin 10002049 | Lupin protein | Not specified in manual | ELISA | Quantitative | ppm | Not specified in manual | 0.2 ppm | 2 ppm | Not specified in manual | Raw to processed foods, from environmental swabs to rinse water | Chickpea 0.0003%, lentil 0.0004%, soy flour (unroasted) 0.07%, soy flour (roasted) 0.0009% and soy lecithin 0.002% | Tested matrices not specified in manual. Ideally would state data for baked goods. | For solid samples: homogenize in mortar or blender, add extraction buffer and suspend whereas for liquid samples extraction buffer is added directly. For samples outside of 6-8 pH, pH should be adjusted with NaOH or HCl. For samples with high tannins or polphenols add 10% fish gelatin | Not stated in manual | Not stated in manual | Not stated in manual | Not stated in manual | Not stated in manual | In a composed / processed food (e.g. corn bread) cross reactivities might be different. Interfering substances (e.g. polyphenols) can be detected by spike experiments | Insufficient data in manual regarding compatible and incompatible matrices
Lupin | - | RBioPharm SureFood ALLERGEN Lupin S3611 | DNA of lupin (Lupinus spp.) according to directive (EC) 1169/2011 | Not specified in manual | PCR | Quantitative | mg/kg | Not specified in manual | ≤ 0.4 mg / kg | 1 mg / kg using SureFood® PREP Advanced, protocol 1 | Corn Flour | Information too sparce to evaluate | None listed | Tested matrices not specified in manual. Ideally would state data for baked goods. | Yes, Advanced Protocol 1 recommended for this matrix | Not stated in manual | Not stated in manual | The technical validation of instruments was performed on Roche LightCycler® 2.0, Roche LightCycler® 480 II, Roche cobas® z 480 Analyzer, Qiagen Rotor-Gene Q, Applied Biosystems 7500, Bio-Rad CFX96, Bio Molecular Systems MIC, LTF MyGo Pro, Agilent AriaDx and Agilent Mx3005P | Not stated in manual | Not stated in manual | Applicable matrices not listed. | Insufficient data in manual regarding compatible and incompatible matrices
Lupin | - | Biocheck Lupin Allergen-Check ELISA R6031/R6030 | Protein(s) not detailed in kit | Yes | ELISA | Quantitative | mg/kg whole lupin seed/flour and mg lupin proteins/kg | Yes, mg/kg whole lupin seed/flour and mg lupin proteins/kg | 0.3 mg whole lupin/kg | 2 mg whole lupin/kg | Not disclosed in manual | Wide range of foods tested, including incurred materials | Chickpea (0.0003%) Lentil (0.0004%). Roasted soya flour (0.0009%), Untreated soya flour (0.07%), Soya Lecithin (0.002%) | None identified, sausage, cookie, fruit juice, ketchup and croquette tested. | No | Not stated in manual | Not stated in manual | None identified | Recovery 99-113% for 5 relevant matrices detailed in manual | None identified | No mention of following AOAC methods for validation. | Specific and sensitive, validation data fairly well documented in manual, tested on relevant incurred matrices.

## Milk
Allergen type | Citing Papers | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target e.g. protein/ peptide/ DNA target. Food or sub-food e.g. whole egg or egg white protein only. | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredients | List relevant/pertinent untested matrices | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data. | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Milk, casein | - | SENSISpec ELISA Casein ppm HU0030027) | Antibody directed against casein | Not clear from manual | ELISA | Quantitative | ppm | Conversion factors: Whole milk 42, Skim milk powder (MoniQA MQA 092014) 4.4, kim milk powder (NIST RM1549) 3.6, Whole milk powder (NIST RM8435) 4.9, Caseinate 1.2 | 0.04 ppm | 0.2 ppm | Not detailed in kit information | cookies, bread crumbs, sausage, orange juice, wine, soy products and chocolate are mentioned in instruction manual | Low (<1.2%) cross-reactivity with ewe's- and goat's milk | Ice cream, soup | Not stated on insert | Not stated on insert | That validation experiments have occurred is referred to in instructions but no further details. | Not stated on insert | Intra-assay precision is 5-11%, inter-assay precision is 8-14% Recovery 80-102% on selected matrices. | Validation data not provided | Full list of tested matrices required
Milk, β-LG | - | Eurofins SENSISpec ELISA ß-Lactoglobulin 96/48 Tests (HU0030002 / HU0030026) | Antibody directed against β-LG | Not clear from manual | ELISA | Quantitative | ppb | Skim milk powder (MoniQA MQA 092014) 4.4, Skim milk powder (NIST RM1549) 3.6, Whole milk powder (NIST RM8435) 4.9, Caseinate 1.2 | 1.5 ppb β-LG | 10 ppb β-LG | Not detailed in kit information | cookies, cereals chocolate, sausage, soy milk, orange juice and white wine | None listed in kit instructions, 41 tested. No significant cross-reactivity with casein, ewe's milk or goat's milk. | Ice cream, soup | Not stated on insert | Not stated on insert | Not stated on insert | Intra-assay precision 7%, inter-assay precision 9-12%. Recovery 70-107% on selected matrices. | Not stated on insert | No cross reactivity with casein or with ewe and goat milk. Named applicable matrices: cookies, cereals, chocolate, sausage, soy milk, orange juice and white wine | Sensitive and specific, some validation information required
Milk, β-LG | - | Biocheck | Protein | No | ELISA | Quantitative | mg skimmed milk allergen/kit, not protein levels | No | <0.1 mg ski9mmed milk allergen /kg allergen, NOT proteinlvels, which are lower | 0.2-0.7 mg skimmed milk allergen /kg allergen, NOT proteinlvels, which are lower | Not detailed in kit information | Not detailed in kit information | Sheep's milk <0.2%, goat's milk <0.002% | Ice cream, soup | Not stated on insert | Not stated on insert | Not stated on insert | Not stated on insert | Not stated on insert | - | Limited list of tested materials provided
Milk, casein | - | Biocheck | Protein | Not clear from manual | ELISA | Quantitative | mg/kg | Not detailed in kit information | <0.2 mg/kg | 0.5-1.0 mg/kg | Not detailed in kit information | Not detailed in kit information | Sheep's milk 1.2%, goat's milk 1.1% | Ice cream, soup | Not stated on insert | Not stated on insert | Not stated on insert | Not stated on insert | Not stated on insert | - | Limited list of tested materials provided
Milk (total) | - | Biocheck | Protein | Not clear from manual | ELISA | Quantitative | mg/kg | Not detailed in kit information | <0.1 mg/kg | 1-27 mg/kg | Not detailed in kit information | Not detailed in kit information | Sheep's milk <0.94%, goat's milk 0.01% | Ice cream, soup | Not stated on insert | Not stated on insert | Not stated on insert | Not stated on insert | Not stated on insert | - | Limited list of tested materials provided
Milk, β-LG | - | Rbiopharm RIDASCREEN® β-Lactoglobulin (Art. No. R4901) | Antibody directed against β-LG' | Yes | ELISA | Quantitative | mg/kg (ppm) | Yes, according to manual, but no according to PTG data (please refer to Objectives 4 and 5). | 2.1 mg/kg | 5 mg/kg | Not detailed in kit information | Beverages containing whey, milk or milk powder as well as in juice, wine, beer and in food like sausages, dressings, bakery products, ice cream, chocolate, yoghurt, etc and hypoallergenic baby food. Manual states that levels may be under-represented for hydrolysed milk. | α-, β-, κ-Casein < 1 % α-Lactalbumin < 1 % bovine serum albumin < 1 % | Soup | Not stated on insert | Not stated on insert | Not stated on insert | Not stated on insert | Not stated on insert | The assay is calibrated to hydrolyzed β-lactoglobulin check samples; the standards consist of native β-lactoglobulin. The assay detects native and processed protein and fragments thereof (during the heating of milk, the native ß-lactoglobulin is partially denatured). | Sensitive and specific, PT data suggests can't inter-convert milk methods, so manufacturer's milk methods are not accurately quantifying as a range of milk methods, ring trial data required
Casein | Gomaa, A. and J. I. Boye (2013). "Impact of thermal processing time and cookie size on the detection of casein, egg, gluten and soy allergens in food." Food Research International 52(1): 483-489., Johnson, P. E., et al. (2014). "A multi-laboratory evaluation of a clinically-validated incurred quality control material for analysis of allergens in food." Food Chemistry 148: 30-36., Ivens, K. O., et al. (2017). "The Effect of Different Methods of Fermentation on the Detection of Milk Protein Residues in Retail Cheese by Enzyme-Linked Immunosorbent Assay (ELISA)." Journal of Food Science 82(11): 2752-2758., Heick, J., et al. (2011). "Application of a Liquid Chromatography Tandem Mass Spectrometry Method for the Simultaneous Detection of Seven Allergenic Foods in Flour and Bread and Comparison of the Method with Commercially Available ELISA Test Kits." Journal of Aoac International 94(4): 1060-1068. | Rbiopharm RIDASCREEN FAST Casein R4612 | Antibodies used specifically detect caseins of cow’s milk | Not clear from manual | ELISA | Quantitative | mg/kg casein | Yes, according to manual, but no according to PTG data (please refer to Objectives 4 and 5). | AEB: 0.12 mg / kg (ppm) casein (mean) 0.07 – 0.19 mg / kg (ppm) casein* RIDA® Extractor 2: 0.71 mg / kg (ppm) casein (mean) 0.41 – 0.95 mg / kg (ppm) casein* *depending on matrix | AEB: 0.5 mg / kg (ppm) casein Extractor 2: 2.5 mg / kg (ppm) casein | Not detailed in kit information | Bakery goods, cake and bread mix, non-hydrolyzed milk-based infant formula, ice cream, beverages, chocolate, wine (more sensitive) and sausages. Extraction for heavily liquid absorbing matrices using RIDA® Extractor 2 (R4613)Preparation of food samples with the RIDA® Extraction Solution (colorless) (R7098) . In case of extremely acidic or basic samples, adjust the sample’s pH value (pH 6.5 - 7.5) to neutral prior to extraction may be necessary. | There is a cross reactivity to sheep’s, goat’s and buffalo’s milk. Any further cross reactivity is not known. Especially, a cross reactivity to ß-lactoglobulin does not exist. | Insufficient data provided regarding matrices tested | Not stated on insert | Not stated on insert | Johnson et al 2014 All 11 kits in ring trial underestimated the level of milk | Not stated on insert | Gomaa and Boye 2013 Thermal processing reduces recovery. Recovery >90% for unbaked cookies. Ivens et al 2017 (1,2) this kit did not respond as well as another when detecting milk in cheese. | Raw fish strongly binds casein- may reduce recovery to 10% (also seen if casein added before cooking fish. In processed foods (e.g. heat treatment, dehydration, etc.), proteins may be altered or fragmented and this may have an impact on the recovery and assay results. In composed / processed food (e.g. maize bread) cross reactivities might be different. Interfering substances (e.g. polyphenols) can be detected by spiking experiments | Sensitive and specific, PT data suggests can't inter-convert milk methods, so manufacturer's milk methods are not accurately quantifying as a range of milk methods, ring trial data required
Milk | Parker, C. H., et al. (2015). "Multi-allergen Quantitation and the Impact of Thermal Treatment in Industry-Processed Baked Goods by ELISA and Liquid Chromatography-Tandem Mass Spectrometry." Journal of Agricultural and Food Chemistry 63(49): 10669-10680. | Rbiopharm FAST Milk R4652 | Antibodies used specifically detect caseins and ß-lactoglobulin of cow’s milk | Not clear from manual | ELISA | Quantitative | mg/kg | Yes, according to manual, but no according to PTG data (please refer to Objectives 4 and 5). | 0.57 mg / kg (ppm) milk protein (mean) 0.3 – 0.8 mg / kg (ppm) milk protein* *depending on matrix | 2.5 mg / kg (ppm) milk protein | Calibrated to skim milk powder | Sausages, ice cream, chocolate, bakery goods, cake and bread mix, soups, sauces, dressings and beverages (juice, wine, beer). | None (checked against 86 matrices) other than Sheep, goat and buffalo milk | Insufficient data provided regarding matrices tested | Certified by the AOAC Performance Tested MethodsSM Program and was assigned PTM Certification No. 101501 by the AOAC Research Institute | Not stated on insert | The in-house validation data was confirmed by an independent lab | Not stated on insert | Parker et al 2015 Recovery reduced with thermal processing. | Composed / processed food (e.g. maize bread) cross reactivities might be different. Interfering substances (e.g. polyphenols) can be detected by spiking experiments. For skim milk powder, whole milk powder or milk, then the ELISA results are quantitative. − Whey (β-lactoglobulin), then the ELISA results may be underestimated. − Caseinates, then the ELISA results may be underestimated | Sensitive and specific, PT data suggests can't inter-convert milk methods, so manufacturer's milk methods are not accurately quantifying as a range of milk methods, ring trial data required
Milk | Johnson, P. E., et al. (2014). "A multi-laboratory evaluation of a clinically-validated incurred quality control material for analysis of allergens in food." Food Chemistry 148: 30-36. | Morinaga Beta-Lactoglobulin M2112 | Beta-Lactoglobulin (whole milk protein) | Not clear from manual | ELISA | Quantitative | ppm | Not detailed in kit information | 0.31 ppm | 0.31-20 ppm (ug protein/ g of food) | Not detailed in kit information | Comprehensive list of ingredients tested | Butter, Carrageenan - see documentation for full list checked | Comprehensive list of ingredients tested | Not stated on insert | Not stated on insert | Johnson et al 2014 All 11 kits in ring trial underestimated the level of milk | Not stated on insert | Not stated on insert | Comprehensive list of ingredients tested, relevent to hyrdolysed products | Sensitive, specific, applies to hydrolysed products
Milk | - | Morinaga Casein M2113 | Casein (whole milk protein) | Not clear from manual | ELISA | Quantitative | ppm | Not detailed in kit information | 0.31 ppm | 0.31-20 ppm (ug protein/ g of food) | Not detailed in kit information | Comprehensive list of ingredients tested | Butter- see documentation for full list checked | Comprehensive list of ingredients tested | Not stated on insert | Not stated on insert | Johnson et al 2014 All 11 kits in ring trial underestimated the level of milk | Not stated on insert | Not stated on insert | Comprehensive list of ingredients tested, relevent to hyrdolysed products | Sensitive, specific, applies to hydrolysed products
Milk | Johnson, P. E., et al. (2014). "A multi-laboratory evaluation of a clinically-validated incurred quality control material for analysis of allergens in food." Food Chemistry 148: 30-36. | RIDASCREEN® FAST β-Lactoglobulin (Art. No. R4912) | Antibody directed against β-LG | Not clear from manual | ELISA | Quantitative | mg/kg (ppm) | Yes, according to manual, but no according to PTG data (please refer to Objectives 4 and 5). | 0.042 mg/kg (0.024-0.073 depending on matrix) | 0.167 mg.kg | Not detailed in kit information | rice crispies, chocolate, and sausage | Sheep, goat and buffalo's milk | None identified | Not stated on insert | Not stated on insert | Johnson et al 2014 All 11 kits in ring trial underestimated the level of milk | Not stated on insert | Not stated on insert | Some validation data missing, more detailed list of compatible matrices required | Sensitive and specific, PT data suggests can't inter-convert milk methods, so manufacturer's milk methods are not accurately quantifying as a range of milk methods, ring trial data required
Milk | - | Biocheck Casein allergen check ELISA | Casein proteins | Yes, mg casein /kg | ELISA | Quantitative | mg casein /kg | Only for measuring casein, not whole milk | 0.05 mg casein/kg | 0.2 mg casein/kg | Not detailed in kit information | Evaluated for a large range of ingredients | Sheep/ewe's milk 1.2%, goat's milk 1.1% | Ice cream, soup | Not stated on insert | Not stated on insert | Not stated on insert | Recoveries 80-102% for 8 matrices tested | Not stated on insert | Some validation data missing | Sensitive, specific, tested against a wide range of matrices
Milk | - | Biocheck Milk allergen check ELISA | Casein and β-LG proteins | mg milk total proteins/kg | ELISA | Quantitative | mg milk total proteins/kg | Apply caution, can only interconvert 'when the origin of the material is known' | 0.1 mg total milk protein/kg | 0.5 mg total milk protein/kg | Not detailed in kit information | Evaluated for a large range of ingredients | Sheep/ewe's milk <0.94%, goat's milk 0.01% | Ice cream, soup | Not stated on insert | Not stated on insert | Not stated on insert | Recoveries 81-104% for 5 matrices for which data is published | Not stated on insert | Some validation data missing | Sensitive, specific, tested against a wide range of matrices
Milk | - | Biocheck β-LG allergen check ELISA | Protein | Yes, µg β-LG/kg | ELISA | Quantitative | mg β-LG/kg | Only for measuring casein, not whole milk | 1.5 µg β-LG/kg | 10 µg β-LG/kg | Not detailed in kit information | Evaluated for a large range of ingredients | Bovine casein <0.02%, ewe's milk <0.2%, goat's milk <0.002% | Ice cream, soup | Not stated on insert | Not stated on insert | Not stated on insert | Recoveries 88-182% for 4 matrices for which data is published | Not stated on insert | Some validation data missing | Sensitive, specific, tested against a wide range of matrices
Milk | Johnson, P. E., et al. (2014). "A multi-laboratory evaluation of a clinically-validated incurred quality control material for analysis of allergens in food." Food Chemistry 148: 30-36. | Romer AgraQuant® β-lactoglobulin 10002034 | Antibody directed against β-LG' | Not clear from manual | ELISA | Quantitative | ppb | Not detailed in kit information | 1.5 ppb | 10 ppb | Not detailed in kit information | raw to processed foods, from environmental swabs to rinse water | Casein <0.02%, ewe’s milk <0.2%, goat’s milk <0.002% | Insufficient data provided regarding matrices tested | Not stated on insert | Not stated on insert | Not stated on insert | Not stated on insert | Not stated on insert | - | Insufficient testing and validation data provided
Milk | Johnson, P. E., et al. (2014). "A multi-laboratory evaluation of a clinically-validated incurred quality control material for analysis of allergens in food." Food Chemistry 148: 30-36. | Romer AgraQuant® Casein 10002037 | Antibody directed against casein | Not clear from manual | ELISA | Quantitative | ppm | Not detailed in kit information | 0.04 ppm | 0.2 ppm | Not detailed in kit information | raw to processed foods, from environmental swabs to rinse water | Ewe’s casein <1.2%, goat’s milk <1.1% and soy protein isolate <0.00004%. These cross-reactivity values were determined with the pure food commodity treated as sample | Insufficient data provided regarding matrices tested | Not stated on insert | Not stated on insert | Johnson et al 2014 All 11 kits in ring trial underestimated the level of milk | Not stated on insert | Not stated on insert | Insufficient data provided regarding matrices tested on | Insufficient testing and validation data provided
Milk | - | Romer AgraQuant® Milk 10002080 | Highly purified polyclonal antibodies raised against milk protein | Not clear from manual | ELISA | Quantitative | ppm | Not detailed in kit information | 0.05 ppm | 0.4 ppm | Not detailed in kit information | raw materials and finished foods and testing trace allergens with rinse waters and enviromental swab | Ewe’s milk 0.94% and goat’s milk 0.01% | Insufficient data provided regarding matrices tested | Not stated on insert | Not stated on insert | Not stated on insert | Not stated on insert | Not stated on insert | Insufficient data provided regarding matrices tested on | Insufficient testing and validation data provided
Milk | Downs, M. L. and S. L. Taylor (2010). "Effects of Thermal Processing on the Enzyme-Linked Immunosorbent Assay (ELISA) Detection of Milk Residues in a Model Food Matrix." Journal of Agricultural and Food Chemistry 58(18): 10085-10091., Gomaa, A. and J. I. Boye (2013). "Impact of thermal processing time and cookie size on the detection of casein, egg, gluten and soy allergens in food." Food Research International 52(1): 483-489., Parker, C. H., et al. (2015). "Multi-allergen Quantitation and the Impact of Thermal Treatment in Industry-Processed Baked Goods by ELISA and Liquid Chromatography-Tandem Mass Spectrometry." Journal of Agricultural and Food Chemistry 63(49): 10669-10680., Ivens, K. O., et al. (2017). "The Effect of Different Methods of Fermentation on the Detection of Milk Protein Residues in Retail Cheese by Enzyme-Linked Immunosorbent Assay (ELISA)." Journal of Food Science 82(11): 2752-2758. | Neogen Veratox for total milk allergen 8470 | Protein | Not clear from manual | ELISA | Quantitative | ppm | Not detailed in kit information | 1 ppm | 2.5 ppm | Non fat dried milk (Test detects casein and whey proteins from cow, goat and sheep) | Juices, cake mixes, cookies, sauces and sorbets, and clean-in-place rinse | Not stated on insert | Not stated on insert | Not stated on insert | Not stated on insert | Not stated on insert | Not stated on insert | Downs and Taylor 2010 Recovery reduced by processing. Parker et al 2015 Recovery reduced with thermal processing. | Concentrated food additives, colors and flavors may cause interferences on ELISA test methods. Neogen Technical Services for validation information. Hydrolyzed and fermented proteins may not be detected using ELISA methods for allergen testing. Although the proteins may be undetectable in the assay because of their properties, there still could be active allergenic protein residue present. | Insufficient testing and validation data provided

## Mollusc
Allergen type | Citing Paper | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method. | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredients | List relevant/pertinent untested matrices | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data. False reporting rate and measurement uncertainty data (inc. precision, recovery). | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Mollusc | - | ELISAKits Mollusc (tropomyosin) ELISA kit, ref. | Protein (tropomyosin) | Yes, measurement of allergenic protein | ELISA | Quantitative | ppb | Details not available | 1.7 ppb mollusc tropomyosin | 10.0 ppb | No relevant information identified online | Soups, fish, bakery, dressings, meat. | No cross reaction was detected for the following: Apricot, Potato, Plum, Bean (white), Chervil, Pine seed, Buckwheat, Chickpea, Pistachio, Cashew, Cherry, Rice, Cayenne, Kiwi, Beef, Chili, Coconut, Bovine gelatine, Cumin, Cress, Rye, Cod, Pumpkin seed, Saccharose, Egg, Caraway, Ewe’s milk, Pea, Lamb, Pork, Egg white powder, Lentil, Celery, Peanut, Lupin, Mustard, Fish gelatine, Corn, Sesame, Barley, Macadamia, Soy, Gliadin, Almond, Soy lecithin, Guar gum, Chestnut, Sunflower seed, Oats, Poppy seed, Tofu, Hazelnut, Nutmeg, Tomato, Isinglass, Bell pepper, Cow’s milk, Chicken, Brazil nut, Walnut, Carob gum, Pecan nut, Wheat, Cocoa, Pepper, Goat’s milk, Carrot, Peach, Onion. | Data too sparce to evaulate | Not specified | Not specified | Not specified | Intra-Assay Precision: 5 – 6%, inter-assay precision: 3%. Recovery: Meat 110%, Fish 103%, Bakery products 100%, vegetable soup 88%, soy sauce 74%. Linearity 89-107%. | No relevant information identified online | More performance data required. | Sensitive method, increaeed availability of validation data needed to evaluate
Mollusc | - | Romer AgraQuant Mollusk, Ref. 10006023 | Protein (tropomyosin) | Yes, measurement of allergenic protein | ELISA | Quantitative | ppb | Details not in manual | 1.7 ppb mollusc tropomyosin | 10.0 ppb | Not detailed in manual | Not detailed in manual | Not detailed in manual | Details too sparce to form an opinion | Not detailed in manual | Not detailed in manual | Not detailed in manual. | None identified | Not detailed in manual | More performance data required. | Sensitive method, increaeed availability of validation data needed to evaluate
Mollusc | - | SureFood Allergen Molluscs, ref. S3613 | DNA. Source of sequence not detailed in kit manual. | No, and the concentration of total DNA detected does not allow a conclusion on the quantity of target DNA. | Real-time PCR | Qualitative | mg/kg | No | ≤0.4 mg/kg | N/A | Not detailed in manual | Not detailed in manual | Not detailed in manual | Not determined. No tested matrices detailed in manual. | Not detailed in manual | Not detailed in manual | Not detailed in manual | Not detailed in manual | No | Manual states, 'spices, herbs, algae, cocoa and further sample matrices might have PCR inhibiting effects'. 50,000 to 60,000 speices of mollusc exist so it important to understand that current methods have not been cross-checked for detection of all species. As squence information is established for each species, theoretical checks regarding applicabilty of the testing method can be conducted, subject to proprietary restrictions. | Does not measure allergenic protein, measures DNA. Not as sensitive as ELISA methods and certain matrices may havinhibitory effects.

## Mustard
Allergen type | Citing papers | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target | Can data be directly applied to the measurement of total food or total protein? | Technology type e.g. ELISA, LC-MS/MS, PCR methods (and specific method type e.g. real-time), LAMP, nanopore, | Qualitative or quantitative scope of method | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredients | List relevant/pertinent untested matrice | Is there more than one extraction method dependant on matrix (e.g. sometimes the case for ELISA)? Need to detail all extraction methods and associated data. | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data(inc. precision and recovery) | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Mustard | - | Romer AgraQuant® Mustard 10002070 | Purified polyclonal antibodies raised against whole mustard | Not specified in manual | ELISA | Quantative | ppm | Not specified in manual | 1 ppm | 2 ppm | Not specified in manual | Raw to processed foods, from environmental swabs to rinse water | Black mustard 50%, brown mustard 59%, charlock mustard 48% and rapeseed 59% | Information too sparce to evaluate | For solid samples: homogenize in mortar or blender, add extraction buffer and suspend whereas for liquid samples extraction buffer is added directly. For samples outside of 6-8 pH, pH should be adjusted with NaOH or HCl. For samples with high tannins or polphenols add 10% fish gelatin | Not specified in manual | Not specified in manual | Not specified in manual | Not specified in manual | Not specified in manual | Cross-reacts with rapeseed. Matrices tested, and whether these were spiked or incurred, not detailed in manual | Matrices tested, and whether these were spiked or incurred, not detailed in manual
Mustard | - | Neogen Veratox for mustard allergen - 8400 | Mustard proteins | The sample should be ground and thoroughly mixed prior to proceeding with the extraction procedure. | ELISA | Quantative | ppm | Not specified in manual | 2.5 ppm | 2.5 ppm | Not stated | Spices, sauces, dressings, meats and clean-in-place rinses | Seeds from the Brassicaceae family (e.g., rapeseed and broccoli seed) | Information too sparce to evaluate | No | Not on insert | Not specified in manual | Not specified in manual | Intra-assay precision 4.9%; inter-assay precision 3.1% - recovery 88% | Not specified in manual | Concentrated food additives, colors and flavors may cause interferences on ELISA test methods. Neogen Technical Services for validation information. Hydrolyzed and fermented proteins may not be detected using ELISA methods for allergen testing. Although the proteins may be undetectable in the assay because of their properties, there still could be active allergenic protein residue present. | Concentrated food additives, colors and flavors may interfere. Hydrolyzed and fermented proteins may not be detected.
Mustard | - | SENSISpec ELISA Mustard 96/48 Tests (HU0030016 / HU0030040) | Mustard proteins | Finely pulverize sample in mortar, impact mill etc | ELISA | Quantative | ppm | Not specified in manual | 1 ppm | 2 ppm | Not stated | Sausage, dressings, soups, cheese and mix herbs | Black Mustard (seed), Brown mustard (seed), Charlock mustard, coriander, rapeseed | Baked goods. | No | Not referenced | Not specified in manual | Not specified in manual | Intra-assay precision 8%; inter-assay precision 12% - recovery 76-98% on selected matricies | Not specified in manual | Cross-reacts with rapeseed | More information regarding applicable amtrices required.
Mustard, celery, Soya | - | RBioPharm SureFood® ALLERGEN 4plex Soya / Celery / Mustard + IAC S3401 | Celery (Apium graveolens), mustard (Brassica carinata, Brassica juncea, Brassica nigra, Sinapis alba, Sinapis arvensis) and soya (Glycine max) DNA sequences according to directive (EC) 1169/2011 | Not specified in manual | PCR | Qualitative | mg/kg | Not specified in manual | ≤ 0.4 mg / kg | Not stated in manual | Corn flour | Information too sparce to evaluate | None stated | Information too sparce to evaluate | No | Not referenced | Not specified in manual | Technical validation of instruments was performed on Roche LightCycler® 480 II, Qiagen Rotor-Gene Q, Agilent Aria Dx, Applied Biosystems 7500, Bio-Rad CFX96, R-Biopharm RIDA®CYCLER and Agilent Mx3005P | Not specified in manual | Not specified in manual | More information regarding applicable matrices required. | More information regarding applicable amtrices required.
Mustard | - | RBioPharm SureFood Allergen Mustard S3609 | DNA of brown (Indian), yellow and black mustard according to directive (EC) 1169/2011 | No | PCR | Quantative | mg/kg | Not specified in manual | ≤ 0.4 mg / kg | 1 mg / kg using SureFood® PREP Advanced, protocol 1 | Corn flour | Information too sparce to evaluate | None stated | Information too sparce to evaluate | Yes, Protocol 1 recommended for this matrix | Not referenced | Not specified in manual | Not specified in manual | Not specified in manual | Not specified in manual | More information regarding applicable amtrices required. | More information regarding applicable amtrices required.

## Peanut
Allergen type | Citing Papers | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target e.g. protein/ peptide/ DNA target. Food or sub-food e.g. whole egg or egg white protein only. Name the target if this information is available e.g. put protein name, gene name or sequence info if available. | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredients | List relevant/pertinent untested matrices | Is there more than one extraction method dependant on matrix (e.g. sometimes the case for ELISA)? Need to detail all extraction methods and associated data. | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data (inc. precision and recovery). | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Peanut | Immer, U., et al. (2004). "RIDASCREEN (R) FAST PEANUT, a rapid and safe tool to determine peanut contamination in food." International Journal of Food Science and Technology 39(8): 869-871., Lacorn, M., et al. (2022). "Validation of the RIDASCREENV (R) Peanut for Determination of Peanut Protein in Cookies, Milk Chocolate, Ice Cream, Trail Mix, Puffed Rice Cereals, and Granola Bar: AOAC Performance Tested Method(SM) 112102." Journal of Aoac International 105(3): 784-801., Trucksess, M. W., et al. (2004). "Preparation of peanut butter suspension for determination of peanuts using enzyme-linked immunoassay kits." Journal of Aoac International 87(2): 424-428., Park, D. L., et al. (2005). "Performance Tested Method multiple laboratory validation study of ELISA-based assays for the detection of peanuts in food." J AOAC Int 88(1): 156-160., Whitaker, T. B., et al. (2005). "Immunochemical analytical methods for the determination of peanut proteins in foods." Journal of Aoac International 88(1): 161-174., Heick, J., et al. (2011). "Application of a Liquid Chromatography Tandem Mass Spectrometry Method for the Simultaneous Detection of Seven Allergenic Foods in Flour and Bread and Comparison of the Method with Commercially Available ELISA Test Kits." Journal of Aoac International 94(4): 1060-1068., Jayasena, S., et al. (2015). "Comparison of Six Commercial ELISA Kits for Their Specificity and Sensitivity in Detecting Different Major Peanut Allergens." Journal of Agricultural and Food Chemistry 63(6): 1849-1855. , Parker, C. H., et al. (2015). "Multi-allergen Quantitation and the Impact of Thermal Treatment in Industry-Processed Baked Goods by ELISA and Liquid Chromatography-Tandem Mass Spectrometry." Journal of Agricultural and Food Chemistry 63(49): 10669-10680., Jayasena, S., et al. (2019). "Comparison of recovery and immunochemical detection of peanut proteins from differentially roasted peanut flour using ELISA." Food Chem 292: 32-38., Park, D. L., et al. (2005). "Performance Tested Method multiple laboratory validation study of ELISA-based assays for the detection of peanuts in food." J AOAC Int 88(1): 156-160., Roland, P., et al. (2005). Inter-laboratory Validation Study of Five Different Commercial ELISA Test Kits for the determination of Peanut Residues in Cookie and Dark Chocolate. | Rbiopharm RIDASCREEN peanut R6811 | Peanut proteins | Yes | ELISA | Quantitative | mg/kg peanut and mg peanut protein/kg | Yes, convert from mg/kg peanut to mg peanut protein/kg | 0.15 mg / kg (ppm) peanut (mean) 0.033 mg / kg (ppm) peanut protein | 0.75 mg / kg (ppm) peanut 0.166 mg / kg (ppm) peanut protein | Not detailed in manual | Puffed rice cereals, granola bar, cookies, ice cream, trail mixes and chocolate | No cross reactivity in 91 commodities | None identified | No | AOAC-RI 112102 validated method | AOAC-RI 112102 validated method | Manual mentions that data 'may be available from inter-laboratory comparisons'. Poms et al 2005 found that this kitand three others all reported false negatives. All kits responded well in the 5-10 mg/kg range but not so well at the 2-2.5 mg/kg level. | Not detailed in manual or on line other than 'available on request'. | Lacorn et al 2022 tested against 91 substances with no cross-reactivity and recovery 99-104%. Trucksess et al 2004 poor recovery compared to one of the other two kits compared. Whitaker et al 2005 Performed favourably on incurred samples | In composed / processed food (e.g. maize bread) cross reactivities might be different. Interfering substances (e.g. polyphenols) can be detected by spiking experiments. In processed foods (e.g. heat treatment, dehydration, etc.), proteins may be altered or fragmented and this may have an impact on the recovery and assay results. In case of extremely acidic or basic samples, adjust the sample’s pH value (pH 6.5 - 7.5) to neutral prior to extraction | Method appears ready for use in incident management.
Peanut | - | Biocheck Peanut Allergen-Check ELISA, Ref. R6021/R6020 | Peanut proteins | mg peanut protein/kg | ELISA | Quantitative | mg protein/kg | Conversion factor between mg/kg peanut and mg/kg peanut protein | <0.5 mg/kg | 1-30 mg peanut protein/kg | Not detailed in manual, although recovery is determined after spiking food matrices | Cookies, breakfast cereal, ice cream, chocolate | Fenugreek (0.0002%), gum arabic (0.003%), soy (0.0002%) (nb no cross reactivity with cumin) | All pertinent FAO-WHO matrices declared | No | Not detailed in manual | Not detailed in manual | Not detailed in manual | Recovery 89-110% for four spiked matrices tested. | Recovery 89-110% for four spiked matrices tested. | No mention of methods used for validation, although relatively high amount of validation included in manual. | Sensitive and specific method. Tested for cross-reactivity against a wide range of foods. Measures peanut protein/kg
Peanut | Jayasena, S., et al. (2015). "Comparison of Six Commercial ELISA Kits for Their Specificity and Sensitivity in Detecting Different Major Peanut Allergens." Journal of Agricultural and Food Chemistry 63(6): 1849-1855., Jayasena, S., et al. (2019). "Comparison of recovery and immunochemical detection of peanut proteins from differentially roasted peanut flour using ELISA." Food Chem 292: 32-38., Daly, M., et al. (2018). "Assessing Almond and Peanut Allergens Using Commercially Available Immunoanalytical Kits and LC-MS/MS: A Case Study." Journal of Aoac International 101(1): 96-101., Sayers, R. L., et al. (2018). "Microfluidic Separation Coupled to Mass Spectrometry for Quantification of Peanut Allergens in a Complex Food Matrix." Journal of Proteome Research 17(1): 647-655. | Romer AgraQuant® Peanut 10001990 | Peanut proteins | Not stated in manual but, theoretically, yes. | ELISA | Quantitative | ppm peanut | Not stated in manual | 0.1ppm | 1ppm | Not stated in manual | Not stated in manual | Fenugreek 0.00015%, Gum arabic 0.003%, soy flour 0.0002% | Data too sparce to evaulate | For solid samples: homogenize in mortar or blender, add extraction buffer and suspend whereas for liquid samples extraction buffer is added directly. For samples outside of 6-8 pH, pH should be adjusted with NaOH or HCl. For samples with high tannins or polphenols add 10% fish gelatin | Not detailed in manual | Not stated in manual | Not stated in manual | Not stated in manual | Jayasena et al 2015 This kit and four others responded less well to processing than one other kit as detect Ara h 3 which is sensitive to processing. Parker et al 2015 Recoveries low for baked incurred goods (<40%) with this ELISA kit and two others compared to an LC-MS method (60-70%). Jayasena et al 2015 This kit performed suitably for minimally processed samples but not as well as another kit as processing increased. | Insufficient data avalable to evaluate applicable matrices. | Kit apprears sufficiently sensitive and specific for peanut.
Peanut | - | Neogen Veratox VIP for Peanut 8430M | Peanut proteins (Ara h 1, Ara h 2, and Ara h 3) | Yes, total peanut and peanut protein | ELISA | Quantitative | ppm | Yes, from total peanut to peanut protein | Not stated | 0.25 ppm peanut protein (1 ppm peanut) | Not stated in manual | Very low levels of raw and further-processed peanut protein in food products, ingredients and clean-in-place (CIP) rinses | Flax seed | Not stated in manual | No | Not stated in insert | No | Park et al 2005 Kit capable of detecting peanut at 5 ppm in spiked samples. Poms et al 2005 found that this kitand three others all reported false negatives. All kits responded well in the 5-10 mg/kg range but not so well at the 2-2.5 mg/kg level. | Not stated in insert | Trucksess et al 2004 Best recovery compared to two other kits. Fu and Maks 2013 Thermal processing reduces recovery. Jayasena et al 2015 This kit performed suitably for minimally processed samples but not as well as another kit as processing increased. | Concentrated food additives, colors and flavors may cause interferences on ELISA test methods. Neogen Technical Services for validation information. Hydrolyzed and fermented proteins may not be detected using ELISA methods for allergen testing. Although the proteins may be undetectable in the assay because of their properties, there still could be active allergenic protein residue present. | Data provided as peanut protein, in line with FAO-WHO guidelines. More data required regarding applicable matrices, but good sensitivity and specificity.
Peanut | Jayasena, S., et al. (2015). "Comparison of Six Commercial ELISA Kits for Their Specificity and Sensitivity in Detecting Different Major Peanut Allergens." Journal of Agricultural and Food Chemistry 63(6): 1849-1855., Parker, C. H., et al. (2015). "Multi-allergen Quantitation and the Impact of Thermal Treatment in Industry-Processed Baked Goods by ELISA and Liquid Chromatography-Tandem Mass Spectrometry." Journal of Agricultural and Food Chemistry 63(49): 10669-10680., Jayasena, S., et al. (2019). "Comparison of recovery and immunochemical detection of peanut proteins from differentially roasted peanut flour using ELISA." Food Chem 292: 32-38., Jayasena, S., et al. (2019). "Improved extraction of peanut residues from a wheat flour matrix for immunochemical detection." Food Chemistry 278: 832-840. | Morinaga Peanut ELISA kit II | Partially purified peanut protein (whole peanut protein) | Yes, total peanut and peanut protein | ELISA | Quantitative | ppm | Yes, from total peanut to peanut protein | 0.31 ppm | 0.31-20 ppm (ug protein/ g of food) | Not stated in manual | Tested against a wide range of matrices | (Some variability between batches) Roasted hazelnut (e.g. 0.32%), octopus (0.24%), millet species (0.3-0.8%), Quinoa (0.21%) - see documentation for full list checked | Not stated in manual | No | Not stated in insert | Guidelines followed not detailed in manual | Not detailed in manual | Not detailed in manual | Jayasena et al 2015 Kit responded best of the six kits as detects Ara h 2 and Ara h6 which are the most potent allergens in peanut and are less susceptible to processing degradation. Parker et al 2015 Recoveries low for baked incurred goods (<40%) with this ELISA kit and two others compared to an LC-MS method (60-70%). Jayasena et al 2015 This kit performed best of six kits for processed samples. Jayasena et al 2019 also reported that the extraction buffer for this kit is more suited to extracting processed protein. | Insufficient data avalable to evaluate applicable matrices. Information regarding the processing state of tested matrices would be beneficial. | Sufficiently sensitive and specific, some low cross-reactivity with hazelnut. Best performance for processed foods as targets Ara h2 and Ara h 6, according to Jayasena et al 2015 and Jayasena et al 2019. Jayasena et al 2019 also reported that the extraction buffer for this kit is more suited to extracting processed protein.
Peanut | - | Morinaga High Sensitive Peanut ELISA kit II | Ara h2, Ara h6 (whole peanut protein) | Yes, total peanut and peanut protein | ELISA | Quantitative | ppm | Yes, from total peanut to peanut protein | 0.2 ppm | 0.2 ppm | Not stated in manual | Tested against a wide range of matrices | (Some variability between batches) Foxtail millet, Common millet, Japanese barnyard millet, Octopus, squid - see documentation for full list checked | Comprehensive list of matrices tested | No, although it is permitted to either do a quick or overnight extraction but this is not sample dependent just convenience | Not stated in insert | Guidelines followed not detailed in manual | Not detailed in manual | Not detailed in manual | Not detailed in manual | Insufficient data avalable to evaluate applicable matrices. Information regarding the processing state of tested matrices would be beneficial. | Sufficiently sensitive and specific..
Peanut | Trucksess, M. W., et al. (2004). "Preparation of peanut butter suspension for determination of peanuts using enzyme-linked immunoassay kits." Journal of Aoac International 87(2): 424-428., Park, D. L., et al. (2005). "Performance Tested Method multiple laboratory validation study of ELISA-based assays for the detection of peanuts in food." J AOAC Int 88(1): 156-160., Whitaker, T. B., et al. (2005). "Immunochemical analytical methods for the determination of peanut proteins in foods." Journal of Aoac International 88(1): 161-174., Fu, T. J. and N. Maks (2013). "Impact of Thermal Processing on ELISA Detection of Peanut Allergens." Journal of Agricultural and Food Chemistry 61(24): 5649-5658., Jayasena, S., et al. (2015). "Comparison of Six Commercial ELISA Kits for Their Specificity and Sensitivity in Detecting Different Major Peanut Allergens." Journal of Agricultural and Food Chemistry 63(6): 1849-1855., Jayasena, S., et al. (2019)., Roland, P., et al. (2005). Inter-laboratory Validation Study of Five Different Commercial ELISA Test Kits for the determination of Peanut Residues in Cookie and Dark Chocolate. "Comparison of recovery and immunochemical detection of peanut proteins from differentially roasted peanut flour using ELISA." Food Chem 292: 32-38., Jayasena, S., et al. (2019). "Improved extraction of peanut residues from a wheat flour matrix for immunochemical detection." Food Chemistry 278: 832-840., Park, D. L., et al. (2005). "Performance Tested Method multiple laboratory validation study of ELISA-based assays for the detection of peanuts in food." J AOAC Int 88(1): 156-160. | Neogen Veratox VIP for Peanut 8430 | Peanut Proteins | Measures ppm peanut protein | ELISA | Quantitative | ppm | Yes, from total peanut to peanut protein | Not stated in manual | 2.5 ppm | Not stated in manual | Cookies, crackers, chocolate bars, ice cream, cereals and in clean-in-place rinses | Not stated in manual | Paste/spread e.g. almond butter | No | Not stated in insert | Yes | Not stated in insert | Intra-assay Precision 4.3%; Inter-assay precision 4.9% recovery 86% | Park et al 2005 successfully detected 60 samples spiked with peanut and 60 blank samples. Whitaker et al 2005 Performed favourably on incurred samples. Jayasena et al 2015 This kit and four others responded less well to processing than one other kit as detect Ara h 3 which is sensitive to processing. | Concentrated food additives, colors and flavors may cause interferences on ELISA test methods. Neogen Technical Services for validation information. Hydrolyzed and fermented proteins may not be detected using ELISA methods for allergen testing. Although the proteins may be undetectable in the assay because of their properties, there still could be active allergenic protein residue present. | Sensitive and measures peanut protein but no cross-reactivity data supplied. Assessed according to AOAC methods by Park et al 2005.
Peanut | - | SENSISpec ELISA Peanut 96/48 Tests (HU0030019, HU0030043) | Peanut proteins | No, measures 'ppm peanut' | ELISA | Quantitative | ppm | Not stated in manual | 0.08 ppm | 1 ppm | LOD Cookies 0.06 ppm, Ice-cream 0.25 ppm, Chocolate 0.24 ppm, Cornflakes 0.25 ppm, Hazelnut 0.14 ppm | Cookies, cereals, ice cream, chocolate and hazelnut (listed in paperwork) | Fenugreek (0.0002%), gum arabic (0.003%), soy (0.0002%) (nb no cross reactivity with cumin) | Paste/spread e.g. almond butter | Not stated | Not referenced | Not applicable | Not stated in insert | Intra-assay Precision 7-10%; Inter-assay precision 2-11% recovery across different matrices 93-102% | Not stated in insert | No conversion factor to convert data to peanut protein supplied. No reference to AOAC methods for validation. | Sensitive, specific.
Peanut | - | SureFood® ALLERGEN Peanut Art. No.: S3603 | DNA of peanut (Arachis hypogaea) according to directive (EC) 1169/2011 | Not stated in manual | PCR | Quantitative | mg/kg | Not stated in manual | ≤ 0.4 mg / kg | 1 mg / kg using SureFood® PREP Advanced, protocol 1 | Corn flour | Not listed in manual | Not listed in manual | Data too sparce to form an opinion | Yes, Protocol 1 recommended for this matrix | Not referenced | Not applicable | Not referenced | Recovery 90-110% for the four matrices detailed. Intra-assay precision 7-10%, inter-assay precision 2-11%. | Not detailed in manual | List of applicable and non-applicable matrices not detailed in manual. | Insufficient data in manual to determine applicability.

## Sesame
Allergen type | Citing papers | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target e.g. protein/ peptide/ DNA target. Food or sub-food e.g. whole egg or egg white protein only. Name the target if this information is available e.g. put protein name, gene name or sequence info if available. | Is method commercially available, or does it have sufficient validation data? | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredients | List relevant/pertinent untested matrices | Is there more than one extraction method dependant on matrix | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data (inc. precision and recovery). | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Sesame | - | Rbiopharm RIDASCREEN FAST sesame R7202 | Protein of sesame | Commercially available | No | ELISA | Quantitative | mg sesame/kg | Not required | 0.14 mg / kg (ppm) sesame (mean) 0.08 – 0.20 mg / kg (ppm) sesame* *depending on matrix | 2.5 mg / kg (ppm) sesame | Not specified in manual | Instant soup, chocolate dessert, bakery products and crackers | None listed | Oil, ready meals, sauces, ice cream | No | Not specified in manual | Not specified in manual | Not specified in manual | Not specified in manual | Not specified in manual | Composed / processed food (e.g. maize bread) cross reactivities might be different. Interfering substances (e.g. polyphenols) can be detected by spike experiments. Adjust the pH to a neutral value for extremely acidic or alkaline samples. Allergen containing samples that have been heat treated show a reduced recovery because the proteins denature and are no longer recognized by the antibody. The reduction in recovery depends strongly on the temperature and the duration of the heat treatment. | Sensitive method, no cross-reacting foods specified in manual.
Sesame | - | Romer AgraQuant® Sesame 10002064 | Highly purified polyclonal antibodies raised against whole sesame | Commercially available | No | ELISA | Quantitative | ppm | Not required | 0.2 ppm | 2 ppm | Not specified in manual | Not specified in manual | Black sesame 30%, cayenne pepper 0.0006%, chia seeds 0.36%, onion 0.0007%, psyllium husk 0.1% and sunflower seeds 0.0003% | Data too sparce to evaluate | For solid samples: homogenize in mortar or blender, add extraction buffer and suspend whereas for liquid samples extraction buffer is added directly. For samples outside of 6-8 pH, pH should be adjusted with NaOH or HCl. For samples with high tannins or polphenols add 10% fish gelatin | No stated on insert | No stated on insert | No stated on insert | No stated on insert | No stated on insert | Insufficient perfomance data in manual. | Sensitive method, no cross-reacting foods specified in manual.
Sesame | - | Morinaga Sesame ELISA Kit II M2121 | 11S globulin (whole sesamie protein) | Commercially available | No | ELISA | Quantitative | ppm | Not required | 0.16 ppm | 0.16 ppm | Not stated in manual | Wide range of processed and unprocessed matrices tested. | Poppy seed (>10%), hemp see (>2.5%) | None identified | No, although it is permitted to either do a quick or overnight extraction but this is not sample dependent just convenience | Not detailed in manual | Not detailed in manual | Not detailed in manual | Not detailed in manual | Not detailed in manual | Cross reacts with poppy seed. | Sensitive method, applicable to processed and unprocessed matrices. Further performance data required.
Sesame | - | Neogen Veratox for Sesame Allergen 8530 | Seasame seed protein | Commercially available | Yes | ELISA | Quantitative | ppm sesame protein | No | 0.16 ppmsesame protein | 2.5 ppm | Not stated in manual | Manual mentions applicability to bakery products, spices, cereal products | almond, macadamia nut, linseed and sunflower seed when tested at 100%. | Ice cream, ready meals/sauces, oil | Sample should be ground and thoroughly mixed prior to proceeding with the extraction procedure. | Not stated on insert | Not stated on insert | Not stated on insert | Not detailed in manual | Not detailed in manual | Cross reacts with almond, macadamia nut, linseed and sunflower seed when tested at 100%. Concentrated food additives, colors and flavors may cause interferences on ELISA test methods. Neogen Technical Services for validation information. Hydrolyzed and fermented proteins may not be detected using ELISA methods for allergen testing. Although the proteins may be undetectable in the assay because of their properties, there still could be active allergenic protein residue present. | Cross reacts with some tree nuts. Further performance and applicabilty data required.
Sesame | - | SENSISpec ELISA Sesame 96/48 Tests HU0030022 / HU0030046 | Sesame proteins | Commercially available | No | ELISA | Quantitative | ppm sesame protein | No | 0.2 ppm | 2 ppm | LOD for soup, ice cream, sausage, salad sauce, cracker | Bakery products, soups, sausage, sauces and ice-cream | Black sesame 30%, cayenne 0.0006%, chia 0.36%, psyllium husk 0.1%, onion 0.0007%, sunflower seeds 0.0003% | Oil | Yes - For solids finely pulverize sample in mortar, impact mill etc, before suspending in extraction and sample dilution buffer; for liquids - taken straight into extraction and sample dilution buffer | Not stated on insert | Not stated on insert | Not stated on insert | Intra-assay precision: 5-12%; Inter-assay precision 4-10%; recovery across matrices 85-109% | Not detailed in manual | Not validated according to AOAC method | Sensitive and specific, applies to processed and unprocessed foods.
Sesame | - | RBioPharm SureFood® ALLERGEN Sesame Art. No.: S3608 | DNA of sesame (Sesamum indicum) according to directive (EC) 1169/2011 | Commercially available | No | PCR | Quantitative | mg/kg | No | ≤ 0.4 mg sesame / kg | 1 mg / kg using SureFood® PREP Advanced, protocol 1 | Not detailed in manual | Not detailed in manual | Cross reactivity was observed with DNA extracts from devil´s claw (Harpagophytum procumbens) with 100% and DNA extracts from Lamiaceae, subfamiliy Nepetoideae with 0.0002 % | Data too sparce to evaluate | No | Not stated on insert | Not stated on insert | The technical validation of instruments was performed Roche LightCycler® 2.0, Roche LightCycler® 480 II, Roche cobas® z 480 Analyzer, Qiagen Rotor-Gene Q, Applied Biosystems 7500, Bio-Rad CFX96, Bio Molecular Systems MIC, LTF MyGo Pro, AriaDx and Agilent Mx3005P | Not stated on insert | Not stated on insert | Does not evaluated protein | Sesnitive. Insufficient performance data. Does not measure the allergenic protein.

## Soybean
Allergen type | Citing Papers | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target e.g. protein/ peptide/ DNA target. Food or sub-food e.g. whole egg or egg white protein only. Name the target if this information is available e.g. put protein name, gene name or sequence info if available. | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredient | List relevant/pertinent untested matrices | Is there more than one extraction method dependant on matrix | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data (inc. precision and recovery). | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Soya | References Binaghi, M. J., et al. (2017). "Soya traces quantification in cookies and extruded products with ELISA kits." Acta Bioquimica Clinica Latinoamericana 51(4): 653-660.,Lacorn, M., et al. (2016). Accurate and Sensitive Quantification of Soy Proteins in Raw and Processed Food by Sandwich ELISA i. X. A. The Pennsylvania State University CiteSeer. | Rbiopharm RIDASCREEN FAST Soya R7102 | Soya proteins | Yes, soya protein | ELISA | Quantitative | mg soya protein/kg | Not applicable, reports soya protein | 0.24 mg / kg (ppm) soya protein (mean) 0.15 – 0.32 mg / kg (ppm) soya protein* *depending on matrix | 2.5 mg / kg (ppm) soya protein | Calibrator: soya protein (soya flour with 39 % soya protein) | Native and heated food samples and beverages. Sausages, dressings, bakery products, ice, chocolate, soup, sauce, margarine | legumes of the tribe Phaseoleae (e.g. beans), the genus Vicia (e.g. fava bean), to dried peas / pea flour and to peanut. Levels of cross-reactivity not stated | None identified | Samples which are solid or liquid without casein are extracted differently, as is food which contains tannin and polyphenol (eg spices, chocolate and coffee) | Not stated in manual | Not stated in manual | Not stated in manual | Not stated in manual | Compared favourably to Romer and Neogen in Binaghi et al. Lacorn et al 2016 reported cross-reactivities in line with those detailed in the manual and these may be relevant from a clinical point of view due to co-sensitivity of these allergens with soyabean. | Composed / processed food (e.g. corn bread) cross reactivities might be different. Interfering substances (e.g. polyphenols) can be detected by spike experiments | Performed well in inter-lab trial.sensitive and specific. Insufficient performance data included in manual. Levels of identified cross-reaqctivity are not stated.
Soya | - | Biocheck Soya Allergen-Check ELISA | Soya trypsin inhibitor | Yes | ELISA | Quantitative | µg soya trypsin inhibitor/kg | Yes, can convert to soya flour (heated or unheated) and heated soya protein. | µg soya trypsin inhibitor/kg | 50-600 µg soya trypsin inhibitor/kg | Not detailed in kit manual, although recovery studies used spiked foods. | Cookie, breakfast cereal, ice cream, chocolate, instant soup and sausage. | Sesame (0.0002%) | Paste, ready meal, but all FAO-WHO-recommended matrices tested. | No | Not detailed in manual | Not detailed in manual | Not detailed in manual | Recovery 77-106% for the six matercies declared. | No | Method used for validation not declared but relatively high amount of validation data included. | Reports mg allergenic allergen/food, sensitive, specific.
Soya | - | Morinaga Soya ELISA kit II M2117 | Beta-conglycinin (whole soy protein) | Yes | ELISA | Quantitative | ppm | Not applicable, reports soya protein | 0.31 ppm | 0.31-20 ppm (ug protein/ g of food) | Not detailed in manual | Tested against a wide range of foods | Beef liver (0.61%), curry powder spice mix (0.38%), Gellan gum. Beans: concord pole, common, kidney (circa 0.4%) and black bean (>20%) - see documentation for full list checked | None identified | No, although it is permitted to either do a quick or overnight extraction but this is not sample dependent just convenience | Not stated in manual | Not stated in manual | Not stated in manual | Not stated in manual | Not stated in manual | Cross reacts with black bean (20%) | Tested on hydrolysed foods, high sensiitivity, wide range of foods tested for cross-reactivity.
Soya | References Binaghi, M. J., et al. (2017). "Soya traces quantification in cookies and extruded products with ELISA kits." Acta Bioquimica Clinica Latinoamericana 51(4): 653-660., | Romer AgraQuant® Soy 10002015 | highly purified polyclonal antibodies raised against soy trypsin inhibitor | Yes | ELISA | Quantitative | ppb | Kit states that conversion factors available to convert the amount of soy into soy product from soy allergen, for soy protein, soy milk, soy flour (roasted and unroated) and tofu. | 16 ppb | 40 ppb | Not detailed in manual | Tested matrices not detailed. | None listed | Data too sparce to evaluate. | For solid samples: homogenize in mortar or blender, add extraction buffer and suspend whereas for liquid samples extraction buffer is added directly. For samples outside of 6-8 pH, pH should be adjusted with NaOH or HCl. For samples with high tannins or polphenols add 10% fish gelatin | Not stated in insert | Not stated in manual | Not stated in manual | Not stated in manual | Only capable of detected soya added to extruded biscuits in Binaghi et al 2017 | Applicable/tested matrices not detailed, nore is validation data. | Insufficient applicabilty and performance data in manual.
Soya | References Binaghi, M. J., et al. (2017). "Soya traces quantification in cookies and extruded products with ELISA kits." Acta Bioquimica Clinica Latinoamericana 51(4): 653-660.,Cucu, T., et al. (2013). "Effect of partial hydrolysis on the hazelnut and soybean protein detectability by ELISA." Food Control 30(2): 497-503., Platteau, C., et al. (2013). "Effect of oxidation in the presence or absence of lipids on hazelnut and soybean protein detectability by commercial ELISA." Food and Agricultural Immunology 24(2): 179-192., Gomaa, A. and J. I. Boye (2013). "Impact of thermal processing time and cookie size on the detection of casein, egg, gluten and soy allergens in food." Food Research International 52(1): 483-489. | Neogen Veratox for soy allergen 8410 | Soy protein (7S conglycinin) | Can report as soy protein, not soy allergenic protein. | ELISA | Quantitative | ppm | Yes, can convert from soy flour to soy protein isolate or soy protein. | 2.5 ppm soy flour | 2.5 ppm soy flour | Not stated in manual | Cookies, crackers, granola/power bars, ice cream, salad dressings, cereals, liquid products (e.g., juices and milk) and clean-in-place rinses. | None | Curries, ready meals | No | Not stated in insert | Not stated in insert | Not stated in insert | Precision data inter-assay precision 2.7%. Intra-assay precision 3.4% recovery mean 86% | Incapable of soya dection in biscuts at 5000 ppm Binaghi et al 2017. Kit showed a lack of robustness with heating (Plateau 2013)and glycation (Plateau et al 2011). Overestimated soyabean level upon partial hydrolysis with pepsin by 260% (Cucu et al 2013). Performed well on raw cookies byt recoveries fell to 33-1% upon baking. | Concentrated food additives, colors and flavors may cause interferences on ELISA test methods. Neogen Technical Services for validation information. Hydrolyzed and fermented proteins may not be detected using ELISA methods for allergen testing. Although the proteins may be undetectable in the assay because of their properties, there still could be active allergenic protein residue present. Insufficient performance data included. | Sensitive method. Insufficient performance data, particularly for cross-reactivity. Complete lack of robustness to processing in literature.
Soya | Instructions available online, https://www.eurofins-technologies.com/sensispec-elisa-soy-96-wells.html | Sensispec ELISA soy HU0030023 / HU0030047 | Soybean trypsin inhibitor (STI, monomeric protein) | Yes | ELISA | Quantitative | ppb soybean trypsin inhibitor | Yes, can convert from soy STI to soy flour (roasted or unroasted), soy protein isolate, soy milk, tofu, textured soy. | 16 ppb | 40 ppb | Not detailed in instructions | Soy flour (unroasted), soy flour (roasted), soy protein isolate, soy milk, textured soy (granulate), tofu, cookies, cereals, chocolate, ice cream, instant soup, sausage. | None listed, 73 non-cross-reacting ingredients lists. No cross-reacting legumes. | None identified | No | None detailed in instruction maunal. | None detailed in instruction maunal. | None detailed in instruction maunal. | Precision data e.g. inter-assay precision 5-13%. Intra-assay precision 6-8% | Not stated in manual | AOAC methods for validation not referenced. | Weakness: conversion factor required when yield results, depending on the type of product (and list of products is limited)
Soya | - | SureFood® ALLERGEN Soya Art. No.: S3601 | soya DNA (Glycine max) according to directive (EC) 1169/2011 | According to manual, report as mg allergenic substance/kg food sample | PCR | Quantitative | mg allergenic substance/kg food sample | Not required, reports mg allergenic substance/kg food sample | ≤ 0.4 mg / kg | 1 mg / kg using SureFood® PREP Advanced, protocol 1 | Corn flour | Not detailed in manual | Not detailed in manual | Data too sparce to evaluate. | Yes, Protocol 1 recommended for this matrix | Not detailed in manual | Not detailed in manual | Internal validation was performed on Roche LightCycler® 480 II, Roche cobas® z 480 Analyzer, Qiagen Rotor-Gene Q, Applied Biosystems 7500, Bio-Rad CFX96, Bio Molecular Systems MIC, LTF MyGo Pro, AriaDx and Agilent Mx3005P. | Not detailed in manual | Not detailed in manual | Performance data, particularly applicability, lacking. | Sensitivity high but data too sperce to evaluate method performance.

## SO2 (refer to official methods)
Allergen type | Citing Papers | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target | Is method commercially available, or does it have sufficient validation data? | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method | Units used in method | Is inter-conversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredients | List relevant/pertinent untested matrices | Is there more than one extraction method dependant on matrix | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data. | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Sulphites | Wilkes, E. (2018). "Sulfur dioxide measurement: Practical measurement of total SO2 in wine." Wine & Viticulture Journal 33(4): 32-34. | AOAC (Method 990.31, 1995) focuses on the use of ion-exclusion chromatography followed by electrochemical (amperometric) detection | Food and beverages | Sufficient validation data | No | Titration | Quantitative | µg SO2/g | No | 10 µg SO2/g | 10 µg SO2/g | Food and beverages | Foods, not dark coloured foods such as caramel colour | Dark coloured foods | Not detailed | No | AOAC | Yes | Yes: sr = 4.7; sR = 8.9; RSDr = 12.3%; RSDR = 21.2% | Not detailed | Not detailed in manual | The method is not applicable to dark coloured foods or ingredients where SO2 is strongly bound e.g. caramel colour. The method does not detect naturally occurring sulphite. | Official method. Not applicable to dark coloured foods or ingredients where SO2 is strongly bound

## Almond
Allergen type | Citing Papers | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target e.g. protein/ peptide/ DNA target. Food or sub-food e.g. whole egg or egg white protein only. Name the target if this information is available e.g. put protein name, gene name or sequence info if available. | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredient | List relevant/pertinent untested matrice | Is there more than one extraction method dependant on matrix | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data (inc. precision, recovery). | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Almond | Röder, M., et al. (2011). "Sensitive and specific detection of potentially allergenic almond (Prunus dulcis) in complex food matrices by Taqman (R) real-time polymerase chain reaction in comparison to commercially available protein-based enzyme-linked immunosorbent assay." Analytica Chimica Acta 685(1): 74-83., Perner, S. P., et al. (2019). "Investigation of Reduced ELISA Recovery of Almond and Hazelnut Traces from Roasted Nut Samples by SDS-PAGE and Mass Spectrometry." Journal of Aoac International 102(5): 1271-1279., Heick, J., et al. (2011). "Application of a Liquid Chromatography Tandem Mass Spectrometry Method for the Simultaneous Detection of Seven Allergenic Foods in Flour and Bread and Comparison of the Method with Commercially Available ELISA Test Kits." Journal of Aoac International 94(4): 1060-1068. | Rbiopharm RIDASCREEN® FAST Almond Art. No.: R6901 | Proteins from almond | No | ELISA | Quantitative | mg almond /kg | Yes, mg almond/kg to mg almond protein/kg, | 0.1 mg / kg (ppm) almond (mean) 0.0 – 0.23 mg / kg (ppm) almond* *depending on matrix | 2.5 – 20 mg / kg (ppm) almond (depending on matrix, level of food processing and extraction method) More sensitive analysis (improved LOQ of 0.625 mg / kg (ppm)) | Not specified in manual | Representative samples for the food groups pastries and sweets: cereals, cookies, ice cream and chocolate. | Cross reaction to apricot stone, mulberries, mahaleb cherry (cross reactions to other plants from genus Prunus are possible).Widespread testing for cross reactivity (80 matrices have been evaluated) | Oil, ready meals | Skimmed milk can either be added to the allergen extraction buffer or the weighed sample | No | No | No | Recovery of spiked samples was between 66 – 124 % (mean: 97.8 %). Other data not included in manual. | Röder et al 2011 reported that this kit showed twice the response of the Neogen Veratox for almond kit. Perner at al 2019 reported that this kit recovered close to 100% in cookies baked at up to 120 °C but almond was not detected above 120 °C suggesting a reduced extractability. Walker et al 2018 reported cross-reactivity with Prunus mahaleb. | Cross-reacts with Prunus species. Processing may affect performance | Sensitive. Much performance data not in manual. Cross-reacts with Prunus species. Not data for oils. Sensitive to thermal processing.
Almond | - | Romer AgraQuant® Almond ELISA test kit 10002021 | Almond protein | No | ELISA | Quantitative | ppm almond | No | 0.2 ppm (0.2 mg/kg) | 0.4 ppm (0.4 mg/kg) | Not specified in manual | Samples can vary from raw to processed foods, from environmental swabs to rinse water | Apricot kernel 100%, cherry kernel 1.7%, peach kernel 16%, plum kernel 1%, mahaleb cherry kernel 1.4%, chia seeds 0.0003% and sweet red pepper seeds 0.00004% | Data too sparce to evaluate. | For solid samples: homogenize in mortar or blender, add extraction buffer and suspend whereas for liquid samples extraction buffer is added directly. Swabbing kit required for swabs For samples outside of 6-8 pH, pH should be adjusted with NaOH or HCl. For samples with high tannins or polphenols add 10% fish gelatin | No | No | No | Not specified in manual | Walker et al 2018 reported cross-reactivity with Prunus mahaleb. | Cross reactivity with Prunus species. Sensitive method. Full performance data not included in manual. | Cross reactivity with Prunus species. No data for oils
Almond | Slotwinski, E., et al. (2018). "Development and Validation of a Quantitative ELISA for the Detection of Almond Residues in Foods." Journal of Aoac International 101(1): 118-123., Röder, M., et al. (2011). "Sensitive and specific detection of potentially allergenic almond (Prunus dulcis) in complex food matrices by Taqman (R) real-time polymerase chain reaction in comparison to commercially available protein-based enzyme-linked immunosorbent assay." Analytica Chimica Acta 685(1): 74-83. | Neogen Vertatox for Almond - 8440 | Almond protein | Total almond | ELISA | Quantitative | ppm total almond | Convert to ppm almond protein | Not stated | 2.5 ppm | Not specified in manual | Cookies, chocolate bars, ice cream, beverages, and clean-in-place rinses | Apple seed and common Prunus genus foods (ie. apricots, cherries, nectarines, peaches, and plums). Cross reactivity has been identified with the seed (stone); however the flesh (seedless fruit) does not cross react. | Oil, ready meals | No | No | No | No | Intra-assay variabliliy 3.5%, inter-assay variability 4.1% recovery mean 83% | Slotwinski et al 2018 seem to have determined kit manual data). Röder et al 2011 reported that this kit showed half the response of the Ridascreen FAST almond kit. | Concentrated food additives, colors and flavors may cause interferences on ELISA test methods. Neogen Technical Services for validation information. Hydrolyzed and fermented proteins may not be detected using ELISA methods for allergen testing. Although the proteins may be undetectable in the assay because of their properties, there still could be active allergenic protein residue present. | Cross reactivity with Prunus species. No data for oils
Almond | - | RBioPharm SureFood® ALLERGEN Almond Art. No.: S3604 | DNA of Almond | No | PCR | Quantitative | mg/kg | No | ≤ 0.4 mg / kg | 1 mg/kg | Corn flour | In general this procedure is applicable for all homogenous food samples. The quantitative determination is inapplicable for swabs, sponges and inhomogeneous fluids/liquids | None | Data too sparce to evaluate. | SureFood® PREP Advanced, protocol 1 recommended | No | No | The technical validation of instruments was performed on Agilent Mx3005P, Agilent AriaDx, Bio-Rad CFX96, Bio-Rad CFX96 Dx, Qiagen Rotor-Gene Q, R-Biopharm RIDA®CYCLER and Roche LightCycler® 480 II | Data not available in manual | Shows only negligible cross-reactivity to Prunoideae foods, compared so significant levels in ELISAs (Röder et al 2011) | Alcohols (e.g. ethanol, isopropanol), surfactants (e.g. CTAB,SDS, Triton X100), salts (e.g. sodium chloride), spices, herbs, algae, cocoa and further sample matrices might have PCR inhibiting effects | Measures DNA rather than the (preferred) protein allergen. Cross reactivity with Prunus species. No data for oils. Shows much less cross-reactivity compared to the ELISAs so may be a valuable second test for incident management.

## Brazil nut
Allergen type | Citing Papers | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredients | List relevant/pertinent untested matrices | Is there more than one extraction method dependant on matrix | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data. | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Brazil Nut | - | Rbiopharm SureFood® ALLERGEN Brazil Nut Art. No.: S3617 | Brazil nut DNA | No | PCR | Quantitative | mg/kg | No | ≤ 0.4 mg / kg | 1 mg / kg | Corn flour | In general this procedure is applicable for all homogenous food samples. The quantitative determination is inapplicable for swabs, sponges and inhomogeneous fluids/liquids | Not detailed in manual | Data too sparce to evaluate | SureFood® PREP Advanced, protocol 1 recommended | No | No | Internal validation was performed on Roche LightCycler® 2.0, Roche LightCycler® 480 II, Roche cobas® z 480 Analyzer, Qiagen Rotor-Gene Q, Applied Biosystems 7500, Bio-Rad CFX96, Bio Molecular Systems MIC, LTF MyGo Pro, AriaDx and Agilent Mx3005P | Not detailed in manual | Not detailed in manual | Alcohols (e.g. ethanol, isopropanol), surfactants (e.g. CTAB,SDS, Triton X100), salts (e.g. sodium chloride), spices, herbs, algae, cocoa and further sample matrices might have PCR inhibiting effects | Measures DNA rather than (preferred) protein. Much performance data lacking.
Brazil Nut | - | Romer | Brazil nut protein assumed | No | Lateral flow device | Qualitative | ppm | No | 5 ppm brazil nut | Not applicable | Not specified in manual | Raw materials, finished food products, rinse waters, environmental samples. Specifics not included. | Not detailed in manual | Data too sparce to evaluate | No | No | No | No | Not detailed in manual | Not detailed in manual | Qualitative method only, no performance data in manual. | Quantitative capability lacking
Pecan | - | Rbiopharm SureFood® ALLERGEN Pecan Art. No.: S3618 | Pecan nut DNA | No | PCR | Quantitative | mg/kg | No | ≤ 0.4 mg / kg | 1 mg / kg | Corn flour | In general this procedure is applicable for all homogenous food samples. The quantitative determination is inapplicable for swabs, sponges and inhomogeneous fluids/liquids | 100% with shagbark hickory (Carya ovata) | Data too sparce to evaluate | SureFood® PREP Advanced, protocol 1 recommended | Not detailed in manual | Not detailed in manual | Internal validation was performed on Roche LightCycler® 2.0, Roche LightCycler® 480 II, Roche cobas® z 480 Analyzer, Qiagen Rotor-Gene Q, Applied Biosystems 7500, Bio-Rad CFX96, Bio Molecular Systems MIC, LTF MyGo Pro, AriaDx and Agilent Mx3005P. | Not detailed in manual | Not detailed in manual | Alcohols (e.g. ethanol, isopropanol), surfactants (e.g. CTAB,SDS, Triton X100), salts (e.g. sodium chloride), spices, herbs, algae, cocoa and further sample matrices might have PCR inhibiting effects | Multiplex method for macadamia, pecan, brazil nut. Only measures DNA not allergenic protein.
Brazil Nut | - | Eurofins SENSISpec ELISA Brazil nut Cat.nr. HU0030018, HU0030042 | Brazil nut protein | No | ELISA | Quantitative | ppm Brazil nut | No | 0.2 ppm | 1 ppm | Not specified in manual | LOD in brackets cookies (0.3 ppm), cereals (0.1 ppm), ice-cream (0.2 ppm), chocolate (0.2 ppm) and sausage (0.2 ppm) | Pecan nut 0.0002% Hazelnut 0.0007% Sesame 0.0003% | None identified, wide range of matrices tested. | Yes - For solids finely pulverize sample in mortar, impact mill etc, before suspending in extraction and sample dilution buffer; for liquids - taken straight into extraction and sample dilution buffer | Not stated | No | No | Intra-assay precision: 8% Inter-assay precision: 2.6-3.8% Recovery ranges 76-112% for different matrices | Not detailed in manual | - | Sensitive, specific method, not validated to AOAC methods.

## Cashew
Allergen type | Citing Papers | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredients | List relevant/pertinent untested matrice | Is there more than one extraction method dependant on matrix | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data. | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Cashew | - | Rbiopharm RIDASCREEN®FAST Cashew Art. No.: R6872 | Cashew proteins | mg cashew /kg | ELISA | Quantitative | mg cashew /kg | No | 0.13 mg / kg (ppm) cashew (mean) 0.10 – 0.19 mg / kg (ppm) cashew* *depending on matrix | 2.5 – 20 mg / kg (ppm) cashew (dependent on matrix, level of food processing and extraction method) | Not detailed in manual | Representative samples for the food groups pastries and sweets: cookies, chocolate and ice cream | Pistachio (level not declared). Widespread testing for cross reactivity (88 matrices have been evaluated) | Data too sparce to evaluate | 4 methods detailed in insert: Sample preparation solid samples (without skim milk powder),Sample preparation liquid samples (without skim milk powder),Sample preparation chocolate with a cocoa fraction of >75 % and pure cocoa(with gelatin allergen extraction buffer),Sample preparation spices (with skim milk Allergen extraction buffer) | Not detailed in manual | Not detailed in manual | Not detailed in manual | Not detailed, although states processing may affect recovery. | Not detailed in manual | Composed / processed food (e.g. corn bread) cross reactivities might be different. Interfering substances (e.g. polyphenols) can be detected by spike experiments, in case of extremely acid or alkaline samples, the pH should be adjusted to a neutral pH | Sensitive method, cross reacts with pistachio.
Cashew | - | Biocheck R6047/R6046 | - | - | ELISA | - | - | - | <0.2 mg/kg | 2-50 mg/kg | - | - | - | - | - | - | - | - | - | - | - | -
Cashew | - | Romer AgraQuant® Cashew 10002094 (COKAL3148) | Cashew proteins | Not detailed in manual | ELISA | Quantitative | ppm | Not detailed in manual | 0.2 ppm cashew | 2 ppm cashew | Not detailed in manual | Samples can vary from raw to processed foods, from environmental swabs to rinse water | Not detailed in manual | Not detailed in manual | For solid samples: homogenize in mortar or blender, add extraction buffer and suspend whereas for liquid samples extraction buffer is added directly. Swabbing kit required for swabs For samples outside of 6-8 pH, pH should be adjusted with NaOH or HCl. For samples with high tannins or polphenols add 10% fish gelatin | Not detailed in manual | Not detailed in manual | Not detailed in manual | Not detailed in manual | Not detailed in manual | Cross-reacts with pistachio, pink peppercorn, hazelnut and walnut. | Sensitive method, cross reacts with pistachio, pink peppercorn, hazelnut and walnut.
Cashew | - | Neogen Veratox® VIP for Cashew 8570 | Cashew proteins | Not detailed in manual | ELISA | Quantitative | ppm cashew protein | Yes, between cashew and cashew protein | 0.03 ppm | 0.2 ppm cashew protein (1 ppm cashew) | - | very low levels of raw and further-processed (including UHT) cashew residues in a wide variety of food products, ingredients, and clean-in-place rinses. This product can also be used to screen environmental swabs. | Kit states 'none shown' although detailed of which foods were tested are not included. | Data too sparce to evaluate | No | Not detailed in manual | Not detailed in manual | Not detailed in manual | Not detailed in manual | Not detailed in manual | Data too sparce to evaluate applicability. | Sensitive method, data too sparce to determine applicability and specificity
Cashew | - | Eurofins, SENSISpec ELISA Cashew 96/48 Tests Cat.nr. HU0030004 / HU0030028 | Cashew proteins | Cashew proteins only | ELISA | Quantitative | ppm | Not detailed in manual | 0.2 ppm | 2 ppm | Not detailed in manual, although recovery study performed in spiked matrices. | LOD in brackets cookies (0.0 ppm), cereals (0.3 ppm), ice-cream (0.3 ppm) and dark chocolate (0.5 ppm) | Pistachio (4%) | Pastes/ready meals/sauces, although FAO-WHO recommended matriced tested. | Yes - For solids finely pulverize sample in mortar, impact mill etc, before suspending in extraction and sample dilution buffer; for liquids - taken straight into extraction and sample dilution buffer | Not detailed in manual | Not detailed in manual | Not detailed in manual | Inter-assay precision: 5-11%, inter-assay precision: 8-14%, recovery across four matrices: 93-109% | Not detailed in manual | Cross-reacts with pistachio | Sensitive, cross-reacts with pistachio.
Cashew | - | Rbiopharm SureFood® ALLERGEN Cashew Art. No.: S3615 | Cashew DNA | No | PCR | Quantitative | mg/kg | No | ≤ 0.4 mg / kg | 1 mg / kg | Corn flour | In general this procedure is applicable for all homogenous food samples. The quantitative determination is inapplicable for swabs, sponges and inhomogeneous fluids/liquids | Not detailed in manual | Data too sparce to evaluate | SureFood® PREP Advanced (S1053), protocol 1 is recommended | Not detailed in manual | Not detailed in manual | The technical validation of instruments was performed on Roche LightCycler® 480 II, Roche cobas® z 480 Analyzer, Qiagen Rotor-Gene Q, Applied Biosystems 7500, Bio-Rad CFX96, R-Biopharm RIDA®CYCLER, AriaDx and Agilent. | Not detailed in manual | Not detailed in manual | alcohols (e.g. ethanol, isopropanol), surfactants (e.g. CTAB,SDS, Triton X100), salts (e.g. sodium chloride), spices, herbs, algae, cocoa and further sample matrices might have PCR inhibiting effects. Very limited detail regarding applicable matrices. | Method does not measure allergen protein. Very limited detail provided regarding applicable matrices.

## Coconut
Allergen type | Citing Papers | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredients | List relevant/pertinent untested matrices | Is there more than one extraction method dependant on matrix | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data. | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Coconut | - | RBiopharm bioavid lateral flow cashew (including hook line), re. BLH710-15 | Protein | No | Lateral Flow Device | Qualitative | Not applicable | Not applicable | 1ppm cashew nuts in water, 10-60 ppm for spiked matrices. | Not applicable | Spiked matrices | Not detailed in manual | Pistachio (10%) | Data too sparce to evaluate | No | No | No | No | Not detailed in manual | No | Lateral flow test only, with 10% cross-reactivity with pistachio. | LFD only, qualitative not quantitative. Hook line included so that strong positive samples do not present as a false negative.
Coconut | - | Romer Agrastrip allergen - coconut | Protein | No | Lateral Flow Device | Qualitative | Not applicable | Not applicable | 10 ppm coconut | Not applicable | Not detailed in manual | Not detailed in manual | Not detailed in manual | Data too sparce to evaluate | No | Not detailed in manual | Not detailed in manual | Not detailed in manual | Not detailed in manual | Not detailed in manual | Lateral flow test only, no cross-reactivity data. Risk of false negatives when high concentration of coconut present. | Risk of false negatives when high concentration of coconut present
Coconut | - | Neogen Veratox® for Coconut 8560 | Coconut proteins | No | ELISA | Quantitative | ppm | No | 1 ppm | Not detailed in manual | Not detailed in manual | Coconut protein residues on surfaces, clean-inplace rinses and in food products such as beverages, cookies, crackers, chocolate bars, ice cream and cereals | Brazil nut, oat, cherry pit, peach pit, cashew, poppy seed, and hazelnut | Oil, soup, ready meals/sauces FAO-WHO ingredients included in testing. | No | Not stated in manual | Not stated in manual | Not stated in manual | Could not access information but manual states 'has undergone rigorous validation process' | Not detailed in manual | Cross-reacts with multiple ingredients. | Sensitive method, understanding of the validation method required for form a full evaluation. Cross-reacts with multiple ingredients.
Coconut | - | Eurofins, SENSISpec ELISA Coconut 96/48 Tests Cat.nr. HU0030005 / HU0030029 | Coconut proteins | No | ELISA | Quantitative | ppm | No | 0.4 ppm | 2 ppm | Not detailed in manual | LODs in brackets cookies (0.4 ppm), cereals (0.2 ppm), ice-cream (0.3 ppm), chocolate (0.5 ppm) and sausage (0.4 ppm) | None detected | Oil, soup, ready meals/sauces FAO-WHO ingredients included in testing. | Yes - For solids finely pulverize sample in mortar, impact mill etc, before suspending in extraction and sample dilution buffer; for liquids - taken straight into extraction and sample dilution buffer | Not stated in manual | Not stated in manual | Not stated in manual | Intra-assay precision: 4%, Inter-assay precision 6-10%, recovery 74-102% | Not stated in manual | Method of validation unknown. | Sensitive, specific method.

## Hazelnut
Allergen type | Citing papers | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredients | List relevant/pertinent untested matrice | Is there more than one extraction method dependant on matrix | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data. | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Hazelnut | Platteau, C., et al. (2013). "Effect of oxidation in the presence or absence of lipids on hazelnut and soybean protein detectability by commercial ELISA." Food and Agricultural Immunology 24(2): 179-192., Piknová, L., et al. (2008). "A novel real-time polymerase chain reaction (PCR) method for the detection of hazelnuts in food." European Food Research and Technology 226(5): 1155-1158., Poms, R. E., et al. (2004). "Methods for allergen analysis in food: a review." Food Additives and Contaminants Part a-Chemistry Analysis Control Exposure & Risk Assessment 21(1): 1-31. | RBbioPharm RIDASCREEN®FAST Hazelnut Art. No.: R6802 | Proteins from hazelnut | Hazelnut protein | ELISA | Quantitative | mg hazelnut protein/kg food | Yes, between hazelnut protein and hazelnut | 0.19 mg / kg (ppm) hazelnut (mean) 0.17 – 0.22 mg / kg (ppm) hazelnut* *depending on matrix | 2.5 – 20 mg / kg (ppm) hazelnut More sensitive analysis (improved LOQ of 0.625 mg / kg (ppm)). | Not declared but reference material was a hazelnut mush with a protein content of 15% | Cereals, bakery products, ice cream and chocolate, various spices | None stated | Oil, ready meals, sauces. FAO-WHO recommended matrices tested. | Skimmed milk can either be added to the allergen extraction buffer or the weighed sample | Collaborative study (14 different laboratories) for hazelnut in dark chocolate resulted in recovery between 93.8 – 118.5 % (mean: 106.3 %). The ELISA has been validated for dark chocolate by the German BVL (§64 method LFGB). | RBiopharm Good ELISA Practice manual followed | Collaborative study (14 different laboratories) for hazelnut in dark chocolate resulted in recovery between 93.8 – 118.5 % (mean: 106.3 %). The ELISA has been validated for dark chocolate by the German BVL (§64 method LFGB). | Not included in manual | Platteau et al 2013 This and three other ELISA kits significantly underestimated hazelnut level in both unprocessed reference materials and in RMs oxidatated with HCl. Piknova et al 2018 reported an LOQ for this kit of 1 mg/kg in incurred materials. Poms et al 2004 reported the same LOQ level. | composed / processed food (e.g. maize bread) cross reactivities might be different. Interfering substances (e.g. polyphenols) can be detected by spike experiments, in case of extremely acid or alkaline samples, the pH should be adjusted to a neutral pH In processed foods (e.g. heat treatment, dehydration, etc.), proteins may be altered or fragmented and this may have an impact on the recovery and assay results. | Sensitive, specific method.
Hazelnut | - | Biocheck R6015/R6014 | Hazelnut protein | Hazelnut protein | ELISA | Quantitative | mg hazelnut protein/kg food | Yes, between hazelnut protein and hazelnut | <0.1 mg hazelnut protein/kg | 1-40 mg/kg | Not declared in manual | Cornflakes, cookies, ice cream, dark chocolate | Walnut (0.0022%) | Data too sparce to evaluate | No | Not specified in manual | No specified in manual | No specified in manual | Recovery ranged from 77-101% for the four matrices specified in the manual. | None specified in manual | Some cross-reactivity with walnut. Methods of validation not disclosed, although relatively high amount of validation data in manual. Applicability to baked goods would be useful data. | No cross-reactivity with a range of foods.Sensitive. Some cross-reactivity with walnut.
Hazelnut | - | Romer AgraQuant® Hazelnut 10002006 (COKAL0348) | Hazelnut proteins | Hazelnut protein | ELISA | Quantitative | ppm | Not specified in manual | 0.3 ppm (0.3 mg/kg) | 1 ppm (1 mg/kg) | Not specifed in manual | Samples can vary from raw to processed foods, from environmental swabs to rinse water | Almond, 0.001%; Walnut, 0.0003% | Data too sparce to evaluate | For solid samples: homogenize in mortar or blender, add extraction buffer and suspend whereas for liquid samples extraction buffer is added directly. Swabbing kit required for swabs For samples outside of 6-8 pH, pH should be adjusted with NaOH or HCl. For samples with high tannins or polphenols add 10% fish gelatin | Not stated | Not stated | Not stated | Not stated | Not stated | Very little validation data provided. Applicability missing. | Sensitive method but insufficient performance data specified to evaluate.
Hazelnut | Platteau, C., et al. (2013). "Effect of oxidation in the presence or absence of lipids on hazelnut and soybean protein detectability by commercial ELISA." Food and Agricultural Immunology 24(2): 179-192. | Neogen Veratox® for Hazelnut 8420 | Hazelnut proteins | Hazelnut protein | ELISA | Quantitative | ppm of total hazelnut | Yes, from total hazelnut to hazelnut protein | 2.5 ppm | 2.5 ppm | Not specifed in manual | Hazelnut protein residues in food products such as cookies, chocolate bars, ice cream and cereal | None stated | Oil, ready meals, sauces. FAO-WHO recommended matrices tested. | Solid sample must be throughly homogenised or sampling may be performed by environmental swabbing | Not specifed in manual | Not specifed in manual | Not specifed in manual | Not specifed in manual | Platteau et al 2013 This and three other ELISA kits significantly underestimated hazelnut level in both unprocessed reference materials and in RMs oxidatated with HCl. | Concentrated food additives, colors and flavors may cause interferences on ELISA test. Hydrolyzed and fermented proteins may not be detected using ELISA methods for allergen testing. Due to the breakdown of proteins to small peptides or amino acids, they may become undetectable by this assay. | Sensitive method. Performance data lacking in manual.
Hazelnut | - | Morinaga Hazelnut ELISA KitII M2119 | Whole hazelnut proteins (11S globulin) | Yes | ELISA | Quantitative | ppm whole hazelnut protein | Theoretically yes, although not detailed in kit | 0.16 ppm [ µg of hazelnut protein / g of food ] Detectable concentration range 0.39 to 25 ppb hazelnut protein | 0.16 ppm [ µg of hazelnut protein / g of food ] | Not specifed in manual | Processed or unprocessed foods extraction solution enables target protein to solubilize even though it was insolubilized by food processing | None stated | Data too sparce to evaluate | No, although it is permitted to either do a quick or overnight extraction but this is not sample dependent just convenience | Not specified in manual | Not specified in manual | Not specified in manual | Not specified in manual | Not specified in manual | Low amount of performance and applicability data provided in manual. | Sensitive method, states that is aplicable to processed and unprocessed foods.
Hazelnut | - | Eurofins, SENSISpec ELISA Hazelnut Cat.nr. HU0030010 / HU0030034 | Hazelnut proteins | No | ELISA | Quantitative | ppm | Not specified in kit | 0.07 ppm | 1 ppm | Cookies, ice cream, sweets, chocolate, food supplement. | LOD in brackets cookies (0.11 ppm), cereals (0.05 ppm), ice cream (0.02 ppm) and chocolate (0.01 ppm) and food supplement (0.04 ppm) | Almond 0.0001% Walnut 0.0003% | Oil, ready meals, sauces. FAO-WHO recommended matrices tested. | No | Not stated | Not stated | Not stated | Intra-assay precision: 4.9%, Inter-assay 2.9%, inter-extraction precision 5.3%, recovery range for six matrices: 77-102% | Not stated | Method of validation not specified but relatively high amount of performance data is provided. | Sensitive and specific method. High number of matrices tested for cross-reactivity and relatively high amount of performance data included in manual.
Hazelnut | Holzhauser, T., et al. (2002). "Detection of potentially allergenic hazelnut (Corylus avellana) residues in food: A comparative study with DNA PCR-ELISA and protein sandwich-ELISA." Journal of Agricultural and Food Chemistry 50(21): 5808-5815., | RBioPharm SureFood® ALLERGEN Hazelnut Art. No.: S3602 | Hazelnut DNA | No | PCR | Quantitative | mg/kg | No | ≤ 0.4 mg / kg | 1 mg/kg | Corn flour | In general this procedure is applicable for all homogenous food samples. The quantitative determination is inapplicable for swabs, sponges and inhomogeneous fluids/liquids | Almond (0.0001%), walnut (0.0003%) | Oil, ready meals, sauces. FAO-WHO recommended matrices tested. | SureFood® PREP Advanced (S1053), protocol 1 is recommended | Not stated | Not stated | Internal validation was performed on Roche LightCycler® 480 II, Roche cobas® z 480 Analyzer, Qiagen Rotor-Gene Q, Applied Biosystems 7500, Bio-Rad CFX96, Bio Molecular Systems MIC, LTF MyGo Pro, AriaDx and Agilent Mx3005P | Recovery 82-102% for six matrices tested | Not specified in manual | Alcohols (e.g. ethanol, isopropanol), surfactants (e.g. CTAB,SDS, Triton X100), salts (e.g. sodium chloride), spices, herbs, algae, cocoa and further sample matrices might have PCR inhibiting effects | Measured only DNA and ot allergenic protein. Sensitive and specific.

## Macadamia
Allergen type | Citing Paper | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredient | List relevant/pertinent untested matrices | Is there more than one extraction method dependant on matrix | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data. | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Macademia | - | Rbiopharm SureFood® ALLERGEN Macadamia Art. No.: S3616 | Macademia nut DNA | No | PCR | Quantitative | mg/kg | No | ≤ 0.4 mg / kg | 1 mg / kg | Corn flour | Manual states, 'In general this procedure is applicable for all homogenous food samples. The quantitative determination is inapplicable for swabs, sponges and inhomogeneous fluids/liquids' but no supporting data included. | Not alluded to in manual | Data too sparce to evaluate | SureFood® PREP Advanced, protocol 1 recommended | Not stated | Not specified in manual | The technical validation of instruments was performed on Roche LightCycler® 2.0, Roche LightCycler® 480 II, Roche cobas® z 480 Analyzer, Qiagen Rotor-Gene Q, Applied Biosystems 7500, Bio-Rad CFX96, Bio Molecular Systems MIC, LTF MyGo Pro, AriaDx and Agilent Mx3005P. | Not specified in manual | Not specified in manual | Alcohols (e.g. ethanol, isopropanol), surfactants (e.g. CTAB,SDS, Triton X100), salts (e.g. sodium chloride), spices, herbs, algae, cocoa and further sample matrices might have PCR inhibiting effects | Measures DNA, not allergenic protein. Very little performance data included in manual.
Macademia | - | Romer AgraQuant Plus macadamia (COKAL1648F) | Macademia nut proteins | No | ELISA | Quantitative | ppm macadamia (raw and finished products) | No | 1 ppm | 1 ppm | Not specified in manual | Data too sparce to evaluate | Not mentioned in manual. Data too sparce to evaluate | Data too sparce to evaluate | No | Not specified in manual | Not specified in manual | Not specified in manual | Not specified in manual | Not specified in manual | Performance and applicability data significantly lacking. | Sensitive. Can not evaluate due to lack of performance data.
Pecan | - | Rbiopharm SureFood® ALLERGEN Pecan Art. No.: S3618 | Pecan nut DNA | No | PCR | Quantitative | mg/kg | No | ≤ 0.4 mg / kg | 1 mg / kg | Corn flour | In general this procedure is applicable for all homogenous food samples. The quantitative determination is inapplicable for swabs, sponges and inhomogeneous fluids/liquids | 100% with shagbark hickory (Carya ovata) | Data too sparce to evaluate | SureFood® PREP Advanced, protocol 1 recommended | Not detailed in manual | Not detailed in manual | Internal validation was performed on Roche LightCycler® 2.0, Roche LightCycler® 480 II, Roche cobas® z 480 Analyzer, Qiagen Rotor-Gene Q, Applied Biosystems 7500, Bio-Rad CFX96, Bio Molecular Systems MIC, LTF MyGo Pro, AriaDx and Agilent Mx3005P. | Not detailed in manual | Not detailed in manual | alcohols (e.g. ethanol, isopropanol), surfactants (e.g. CTAB,SDS, Triton X100), salts (e.g. sodium chloride), spices, herbs, algae, cocoa and further sample matrices might have PCR inhibiting effects | Multiplex method for macadamia, pecan, brazil nut. Only measures DNA not allergenic protein.
Macademia | - | Eurofins, SENSISpec ELISA Macadamia nut, Cat.nr. HU0030013 / HU0030037 | Macademia nut proteins | No | ELISA | Quantitative | ppm macadamia nut | No | 0.1 ppm | 1 ppm | Not specified in manual | LOD in brackets cookies (0.2 ppm), cereals cornflakes (0.1 ppm), ice-cream (0.2 ppm), chocolate and sausage | Hazelnut (0.0002%), Walnut 0.0010% | Oil, ready meals/sauces, FAO-WHO matrices have been tested. | Yes - For solids finely pulverize sample in mortar, impact mill etc, before suspending in extraction and sample dilution buffer; for liquids - taken straight into extraction and sample dilution buffer. | Not stated | Not specified in manual | Not specified in manual | Intra-assay precision: 2-5% Inter-assay precision: 3-6% Recovery ranges 83-106% for different matrices | Not specified in manual | Validation method not specified, although relatively high amount of validation data included in manual. | Sensitive, applies to processed and raw foods.

## Pecan
Allergen type | Citing Papers | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredients | List relevant/pertinent untested matrices | Is there more than one extraction method dependant on matrix | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data. | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Pecan | - | Rbiopharm SureFood® ALLERGEN Pecan Art. No.: S3618 | Pecan nut DNA | No | PCR | Quantitative | mg/kg | No | ≤ 0.4 mg / kg | 1 mg / kg | Corn flour | In general this procedure is applicable for all homogenous food samples. The quantitative determination is inapplicable for swabs, sponges and inhomogeneous fluids/liquids | 100% with shagbark hickory (Carya ovata) | Data too sparce to evaluate | SureFood® PREP Advanced, protocol 1 recommended | Not detailed in manual | Not detailed in manual | Internal validation was performed on Roche LightCycler® 2.0, Roche LightCycler® 480 II, Roche cobas® z 480 Analyzer, Qiagen Rotor-Gene Q, Applied Biosystems 7500, Bio-Rad CFX96, Bio Molecular Systems MIC, LTF MyGo Pro, AriaDx and Agilent Mx3005P. | Not detailed in manual | Not detailed in manual | Alcohols (e.g. ethanol, isopropanol), surfactants (e.g. CTAB,SDS, Triton X100), salts (e.g. sodium chloride), spices, herbs, algae, cocoa and further sample matrices might have PCR inhibiting effects. | Multiplex method for macadamia, pecan, brazil nut. Only measures DNA not allergenic protein.
Pecan | - | Biocheck R6043/R6042 | Pecan nut DNA | No | ELISA | Quantitative | mg pecan/kg | Yes, pecan to pecan protein | <0.2 mg/kg | 2-60 mg/kg | Not specified in manual | Tested against a long list of (mostly raw) ingredients. | Hazelnut (0.0003%), carrot (0.006%), macadamia nut (0.0003%), walnut (0.0035%) | Oil, sauce/ready meals, but WHO-FAO matrices tested. | No | Not specified in manual | Not specified in manual | Not specified in manual | Not specified in manual | Not specified in manual | Validation method not detailed, althought relatively high amount of perfomrnace data included. | Sensitive, applicable to relevent matrices.
Pecan | - | Eurofins, SENSISpec ELISA Pecan nut, Cat.nr. HU0030020 / HU0030044 | Pecan nut proteins | No | ELISA | Quantitative | ppm pecan nut | No | 0.2 ppm | 2 ppm | Not specified in manual | LOD in brackets cookies (0.2 ppm), cereals cornflakes (0.2 ppm), ice-cream (0.1 ppm), chocolate (0.4 ppm) and sausage (0.1 ppm) | Hazelnut 0.0003% Carrot 0.0006% Macadamia nut 0.0003% Walnut 0.0035% | Oil, sauce/ready meals, but WHO-FAO matrices tested. | Yes - For solids finely pulverize sample in mortar, impact mill etc, before suspending in extraction and sample dilution buffer; for liquids - taken straight into extraction and sample dilution buffer | Not specified in manual | Not specified in manual | Not specified in manual | Intra-assay precision: 8% Inter-assay precision: 6-11% Recovery ranges 78-114% for different matrices | Not specified in manual | Validation method not detailed, althought relatively high amount of perfomrnace data included. | Sensitive, applicable to relevent matrices. Relatively high amount of performnace data available in manual.

## Pistachio
Allergen type | Citing Papers | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredients | List relevant/pertinent untested matrices | Is there more than one extraction method dependant on matrix | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data. | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Pistachio | - | Rbiopharm SureFood® ALLERGEN Pistachio Art. No.: S3614 | Pistachio DNA | No | PCR | Quantitative | mg pistachio/kg | No | ≤ 0.4 mg / kg | 1 mg / kg | Corn flour | In general this procedure is applicable for all homogenous food samples. The quantitative determination is inapplicable for swabs, sponges and inhomogeneous fluids/liquids | Not specified in manual | Data too sparce to evaluate | SureFood® PREP Advanced, protocol 1 recommended method | Not specified in manual | Not specified in manual | The technical validation of instruments was performed on Roche LightCycler® 2.0, Roche LightCycler® 480 II, Roche cobas® z 480 Analyzer, Qiagen Rotor-Gene Q, Applied Biosystems 7500, Bio-Rad CFX96, Bio Molecular Systems MIC, LTF MyGo Pro, AriaDx and Agilent Mx3005P. | Not specified in manual | Not specified in manual | Performance data not in manual to assess applicability | Sensitive. Measures only levels of pistachio DNA and not the allergenic protein.
Pistachio | - | Biocheck Pistachio Allergen-Check ELISA R6043/R6042 | Pistachio proteins | No | ELISA | Quantitative | mg pistachio/kg | Yes from pistachio to pistachio protein | <0.2 mg/kg | 1-25 mg/kg | Not specified in manual | Cookies, cereal, ice cream, chocolate, sausage | Cross-reacts with cashew nut (<0.0025%), chia, pumpkin seed, hazelnut and poppy seed (<0.0002%), chilli (<0.0005%), cayenne (<0.0004%), pepper (<0.0025%), Quinoa (<0.002%) | Oil, ready meals/sauces | No | Not specified in manual | Not specified in manual | Not specified in manual | Recovery 81-96% for five spiked matrices teasted. | Not specified in manual | Method of validation not specified. Multiple low levels of cross reactivity. | Sensisitive, relatively high amount of performance data in manual. PCR methods may offer higher specificity (Sanchiz et al 2017).
Pistachio | Sanchiz, A., et al. (2017). "Detection of pistachio allergen coding sequences in food products: A comparison of two real time PCR approaches." Food Control 75: 262-270. | Romer AgraQuant® Pistachio 10002086 (COKAL2748) | Pistachio proteins | No | ELISA | Quantitative | mg/kg | No | 0.13 ppm | 1 ppm | Not specified in manual | Samples can vary from raw to processed foods, from environmental swabs to rinse water | Cashew <0.0025%, cayenne, <0.0004%, chia <0.0002%, chili <0.0005%, hazelnut <0.0002%, pumpkin seed <0.0002%, poppy seed <0.0002%, pepper <0.0025% | Tested foods are not detailed | For solid samples: homogenize in mortar or blender, add extraction buffer and suspend whereas for liquid samples extraction buffer is added directly. Swabbing kit required for swabs For samples outside of 6-8 pH, pH should be adjusted with NaOH or HCl. For samples with high tannins or polphenols add 10% fish gelatin | Not specified in manual | Not specified in manual | Not specified in manual | Not specified in manual | Not specified in manual | Applicable matrices not listed in detail. | Sensitive method, multiple cross-reactivities. PCR methods may offer higher specificity (Sanchiz et al 2017)..
Pistachio | - | Eurofins, SENSISpec ELISA Pistachio, Cat.nr. HU0030021, HU0030045 | Pistachio proteins | No | ELISA | Quantitative | ppm pistachio | No | 0.07 ppm | 1 ppm | Not specified in manual | LOD in bracket cookies (0.18 ppm), ice cream (0.06 ppm), sweets (0.02 ppm), chocolate (0.15 ppm), spices (0.46 ppm) and food supplements (0.13 ppm) | Poppy seed 0.0001% Rapeseed 0.0002% Sesame 0.0001% Horseradish 0.0001% Flaxseed 0.0001% Cashew 0.0033% Mustard, yellow 0.0002% | Oil, ready meals/sauces. WHO-FAO recommended matrices have been tested. | Yes - For solids finely pulverize sample in mortar, impact mill etc, before suspending in extraction and sample dilution buffer; for liquids - taken straight into extraction and sample dilution buffer; samples containing polyphenol add PVP buffer; rinse water samples adjust pH and dilute in less buffer than for liquid; Swab samples, swab is moistened and after swabbing in placed in 1 mL of liquid in a test tube | Not specified in manual | Not specified in manual | Not specified in manual | Intra-assay precision: 6.4% Inter-assay precision: 7.8% Inter-Extraction Precision: 5.8% Recovery ranges 72-90% for different matrices | Sanchiz et al 2017 reported 2 false positives out of 14 samples, suggesting cross-reactivity with hazelnut and cashew. | Cross reacts with multiple foods, validation method not stipulated although relatively high amount of performance data included in manual. | Sensitive. Cross reacts with multiple foods. PCR methods may offer higher specificity (Sanchiz et al 2017).

## Walnut
Allergen type | Citing Paper | Method name, manufacturer name and manufacturer reference number (or similar) / author reference. | Analytical target | Can data be directly applied to the measurement of total food or total protein? | Technology type | Qualitative or quantitative scope of method | Units used in method | Is interconversion of units between methods applicable? | Limit of Detection (LOD), including units | Limit of Quantitation (LOQ), including units | Matrix/matrices used for determination of LOD or LOQ | Specificity: compatible matrices evaluated? | List cross-reacting ingredients | List relevant/pertinent untested matrices | Is there more than one extraction method dependant on matrix | Quality Assurance data and grade, level of accreditation (e.g. UKAS), validation, standardisation (e.g. BSI, CEN, AOAC, ISO). | Have AOAC guidelines for standardized calculations of LOD and LOQ been followed where relevant? | Ring trial data availability. Are ring trials available? Is ring trial data published in the manufacturer's information? | False reporting rate and measurement uncertainty data. | Are any cross-evaluation of methods data available? | Gaps identified | Overall summary of method e.g. strengths and weaknesses of method, evidence gaps.
Walnut | Linacero, R., et al. (2016). "Detection by real time PCR of walnut allergen coding sequences in processed foods." Food Chemistry 202: 334-340. | Romer AgraQuant® Walnut 10002030 (COKAL0948) | Walnut proteins | No | ELISA | Quantitative | ppm | Yes, whole walnut to walnut protein | 0.35 ppm | 2 ppm | Corn flour | Samples can vary from raw to processed foods, from environmental swabs to rinse water | Cashew 0.0002%, chicken 0.0002% and pecan nut 0.0005% | Data too sparce to evaluate | For solid samples: homogenize in mortar or blender, add extraction buffer and suspend whereas for liquid samples extraction buffer is added directly. Swabbing kit required for swabs For samples outside of 6-8 pH, pH should be adjusted with NaOH or HCl. For samples with high tannins or polphenols add 10% fish gelatin | Not specified in manual | Not specified in manual | Not specified in manual | Performance data lacking in manual | Linacero et al 2016 report good sensitivity in a qualitative study. | Data on applicable matrices and performance not included in manual. | Sensitive method, some cross-reactivity reported.
Walnut | - | Biocheck Walnut Allergen-Check R6017/R6016 | Walnut proteins | No | ELISA | Quantitative | mg walnut/kg | Yes, whole walnut to walnut protein | <0.6 mg walnut/kg | 2-50 mg/kg | Not specified in manual | Cookies, breakfast cereal, ice cream, chocolate | Cashew 0.0002%, chicken 0.0002% and pecan nut 0.0005% | Oil, sauce/ready meals | No | Not specified in manual | Not specified in manual | Not specified in manual | No precision data in manual. Recovery 87-106% on the four matrices tested. | Not specified in manual | Validation method not declared, although relatively high amount of performnace data declared in manual | Sensitive method, some cross-reactivity reported.
Walnut | - | RBioPharm SureFood® ALLERGEN Walnut Art. No.: S3607 | Walnut DNA | No | PCR | Quantitative | mg/kg | No | ≤ 0.4 mg / kg | 1 mg / kg | Corn flour | In general this procedure is applicable for all homogenous food samples. The quantitative determination is inapplicable for swabs, sponges and inhomogeneous fluids/liquids. | Not detailed in manual | Data too sparce to evaluate | SureFood® PREP Advanced, protocol 1 recommended | Not specified in manual | Not specified in manual | Internal validation was performed on Roche LightCycler® 2.0, Roche LightCycler® 480 II, Roche cobas® z 480 Analyzer, Qiagen Rotor-Gene Q, Applied Biosystems 7500, Bio-Rad CFX96, Bio Molecular Systems MIC, LTF MyGo Pro, AriaDx and Agilent Mx3005P. | Not specified in manual | Not specified in manual | Insufficient data to evaluate applicability or performance. Alcohols (e.g. ethanol, isopropanol), surfactants (e.g. CTAB,SDS, Triton X100), salts (e.g. sodium chloride), spices, herbs, algae, cocoa and further sample matrices might have PCR inhibiting effects | Sensitive method. Applicable matrices not clear. Measures only DNA and not the level of allergenic protein.
Walnut | Vencia, W., et al. (2019). "Effects of thermal treatment on walnut detection and allergenicity." Journal of the Science of Food and Agriculture 99(5): 2636-2640. | Neogen BioKits Walnut Assay Kit 902085J | Walnut proteins | Walnut protein | ELISA | Quantitative | ppm walnut | Yes, walnut to walnut protein | 0.25 ppm | 2.4 ppm | Not specified in manual | Walnut in raw and processed foods and environmental swabs | Manual states, 'Of a large panel of commodities including nuts, pulses, grains, proteins and seeds the following were reactive: pecan (2.3%); quinoa (0.0012%); pistachio (0.0009%); hazelnut (0.0005%); buckwheat (0.00024%)' | Data too sparce to evaluate | Solid sample must be throughly homogenised or sampling may be performed by environmental swabbing | Not specified in manual | Not specified in manual | Not specified in manual | Not specified in manual | Heick et al reported an overestimation (530%) in raw flour and a significant underestimation (40%) in baked bread. Vencia et al 2019 reported reduced recovery with thermal processing and concluded that kit may not be suitable for highly processed products. | Applicability and other performance data lacking. | Sensitive method, some cross-reactivity reported. Several cross-reactivities, performance data lacking. More recovery data required for processed samples.
Walnut | - | Eurofins, SENSISpec ELISA Walnut, Cat.nr. HU0030024 / HU0030048 | Walnut proteins | Walnut protein | ELISA | Quantitative | ppm | Yes, walnut to walnut protein | 0.3 ppm | 2 ppm | Not specified in manual | LOD not given cookies, cereals, ice cream and chocolate | Hazelnut 0.0006% Pecan 3.5% | Data too sparce to evaluate | No | Not specified in manual | Not specified in manual | Not specified in manual | Intra-assay precision: 4-9% Inter-assay precision: 12% Recovery ranges 60-106% for different matrices | Not specified in manual | Cross reacts with pecan (3.5%) and hazelnut. | Sensitive method, cross-reacts with pecan (3.5%)and hazelnut.