## ELISA
Appendix 3 Section 8 Table 1. Costs to set up a testing laboratory - ELISA
ELISA | Cost, £ | Analyst time requirement, days
Set up costs
Plate reader and associated software | 15967 | -
Optional plate washer | 14261 | -
Price per batch of 16 samples in duplicate, plus QCs
Positive Reference material, per assay | 13 | -
Negative reference material, per assay | 13 | -
Test kit to analyse for 16 analyses in duplicate, on average plus positive and negative QCs in duplicate. | 250 | -
Staff time for extraction and analysis of 16 samples in duplicate, plus an over-spiked matrix match sample, QCs/reference material(non-automated analysis) | - | 1
In-house validation costs
Users may opt to accept some validation data, such as specificity and LOQ, from the parameters supplied with the kit. In-house validation testing may include:
Inter-assay precision, per allergen: | - | 3
Intra-assay precision, per allergen: | - | 3
Applicabilty, per allergen: | - | 3
Sensitivity (LOD), per allergen, in a range of matrices: | - | 1
Measurement uncertainty | - | 0.5
Repeatability: 3 days | - | 3
Cost of 4x 48-well kits required for validation | 1000 | -
Participation in an average of two proficiency testing rounds per year | 500 | -
Total direct costs, £ | 32004 | -
Total time, days | - | 14.5
Comment: | Single allergen tested per test/per validation

## PCR
Appendix 3 Section 8 Table 1. Costs to set up a testing laboratory - PCR
PCR | Cost, £ | Analyst time requirement, days
Set up costs
Real-time PCR instrument | 61350 | -
Laminar flow hood with UV irradiation | 6765 | -
Price per batch of 40 samples in duplicate, plus QCs
Positive Reference material, per assay | 13 | -
Negative reference material, per assay | 13 | -
Test kit to analyse 40 samples in duplicate plus QCs in dupluicate. | 664 | -
Staff time for extraction and analysis of 20 samples in duplicate, plus an over-spiked matrix match sample, QCs/reference material(non-automated analysis) | - | 1
In-house validation costs
Users may opt to accept some validation data, such as specificity and LOQ, from the parameters supplied with the kit. In-house validation testing may include: | - | -
Inter-assay precision, per allergen: | - | 3
Intra-assay precision, per allergen: | - | 3
Applicabilty, per allergen: | - | 3
Sensitivity (LOD), per allergen, in a range of matrices: | - | 1
Measurement uncertainty | - | 0.5
Repeatability | - | 3
Cost of 2x kits required for validation | 1328 | -
Participation in an average of two proficiency testing rounds per year | 500 | -
Total direct costs, £ | 70633 | -
Total time, days | - | 14.5
Comment | Single allergen tested per test/per validation, with current test kits available

## LC-MS
Appendix 3 Section 8 Table 1. Costs to set up a testing laboratory - LC-MS
LC-MS | Cost, £ | Analyst time requirement, days
Set up costs
LC-MS triple quadrupole (MRM) instrument and associated software | 415940 | -
Price per batch of 16 samples in duplicate, plus QCs
Positive Reference material/calibration standards, per batch | 13 | -
Negative reference material/calibration standards per batch | 13 | -
Consumables to analyse 40 samples in duplicate plus QCs in dupluicate. | 50 | -
Staff time for extraction, digestion, clean-up and analysis | - | 2.2
In-house validation costs
Users may opt to accept some validation data, such as specificity and LOQ, from the parameters supplied with the kit. In-house validation testing may include: | - | -
Inter-assay precision, per allergen: | - | 1
Intra-assay precision, per allergen: | - | 1.5
Applicabilty, per allergen: | - | 1.5
Sensitivity (LOD), per allergen, in a range of matrices: | - | 1
Measurement uncertainty | - | 0.5
Repeatability: 3 days | - | 4.5
Cost of solvent consumables required for validation | 227 | -
Participation in an average of two proficiency testing rounds per year | 500 | -
Total direct costs, £ | 416743 | -
Total time, days | - | 12.2
Comment | Multiple (e.g. 5-6) allergens tested per test/per validation